## Rs
BR2330				BR1632				BR2325				BR2477				BR2032				BR1225
Capacity	255	mA-hr		Capacity	125	mA-hr		Capacity	160	mA-hr		Capacity	1000	mA-hr		Capacity	190	mA-hr		Capacity	48	mA-hr
Diameter	23	mm		Diameter	16	mm		Diameter	23	mm		Diameter	24	mm		Diameter	20	mm		Diameter	12	mm
Voltage	2.7	V		Voltage	2.7	V		Voltage	2.7	V		Voltage	2.7	V		Voltage	2.7	V		Voltage	2.7	V
Thickness	3.0	mm		Thickness	3.2	mm		Thickness	2.5	mm		Thickness	7.7	mm		Thickness	3.2	mm		Thickness	2.5	mm

	Ten-Ohm	Rs			Ten-Ohm	Rs			Ten-Ohm	Rs			Ten-Ohm	Rs			Ten-Ohm	Rs			Ten-Ohm	Rs
Number	(V)	(Ohm)		Number	(V)	(Ohm)		Number	(V)	(Ohm)		Number	(V)	(Ohm)		Number	(V)	(Ohm)		Number	(V)	(Ohm)
1	0.781	24.6		1	0.462	48.4		1	0.960	18.1		1	0.970	17.8		1	0.791	24.1		1	0.534	40.6
2	0.734	26.8		2	0.430	52.8		2	0.960	18.1		2	1.033	16.1		2	0.789	24.2		2	0.534	40.6
3	0.757	25.7		3	0.462	48.4		3	0.972	17.8		3	0.998	17.1		3	0.779	24.7		3	0.504	43.6
4	0.748	26.1		4	0.446	50.5		4	0.958	18.2		4	0.966	18.0		4	0.776	24.8		4	0.484	45.8
5	0.751	26.0		5	0.464	48.2		5	0.973	17.7		5	0.876	20.8		5	0.782	24.5		5	0.439	51.5
6	0.742	26.4		6	0.452	49.7		6	1.010	16.7		stdev	0.058	1.8		6	0.775	24.8		6	0.524	41.5
7	0.835	22.3		7	0.447	50.4		stdev	0.020	0.5		ave	0.969	18.0		7	0.784	24.4		7	0.507	43.3
8	0.745	26.2		8	0.453	49.6		ave	0.972	17.8						8	0.766	25.2		8	0.502	43.8
9	0.693	29.0		9	0.458	49.0										stdev	0.008	0.4		9	0.468	47.7
10	0.792	24.1		10	0.450	50.0										ave	0.780	24.6		stdev	0.032	3.6
11	0.759	25.6		stdev	0.010	1.4														ave	0.500	44.2
12	0.755	25.8		ave	0.452	49.7
13	0.722	27.4
14	0.764	25.3
15	0.756	25.7
16	0.704	28.4
17	0.752	25.9
18	0.762	25.4
19	0.720	27.5
20	0.715	27.8
stdev	0.032	1.5
ave	0.749	26.1

## Matching
Tank Tuning: R1, L1, C1, R3, C2 || L2 || R1

Component	Unit	A	B	C	D	E		Component	Unit	A	B	C	D	E
R1	Ohm	8000	8000	8000	8000	8000		R1	Ohm	8.00E+003	8.00E+003	8.00E+003	8.00E+003	8.00E+003
R2	Ohm	50	100	50	100	50		R2	Ohm	5.00E+001	1.00E+002	5.00E+001	1.00E+002	5.00E+001
R3	Ohm	0	0	0	0	1000		R3	Ohm	0.00E+000	0.00E+000	0.00E+000	0.00E+000	1.00E+003
C1	pF	10000	1.3	1	10000	10000		C1	F	1.00E-008	1.30E-012	1.00E-012	1.00E-008	1.00E-008
C2	pF	5	10	9	4	11		C2	F	5.00E-012	1.00E-011	9.00E-012	4.00E-012	1.10E-011
L1	nH	150	300	150	300	0		L1	H	1.50E-007	3.00E-007	1.50E-007	3.00E-007	0.00E+000
L2	nH	10000	100	100	10000	100		L2	H	1.00E-005	1.00E-007	1.00E-007	1.00E-005	1.00E-007

		Gain Magnitude					Gain Phase
f (MHz)	w (rad/s)	A	B	C	D	E				A	B	C	D	E
10.0	62831853	0.982	0.001	0.000	0.958	0.006				3.54	-0.51	-0.23	7.26	89.70
10.2	64088490	0.982	0.001	0.000	0.958	0.006				3.43	-0.52	-0.23	7.06	89.69
10.4	65370260	0.982	0.001	0.000	0.959	0.006				3.33	-0.54	-0.23	6.86	89.68
10.6 / Frequency (MHz) / Gain (V/V)	66677665	0.983	0.001	0.000	0.959	0.006				3.22	-0.55	-0.24	6.67	89.67
10.8	68011218	0.983	0.001	0.000	0.960	0.007				3.12	-0.56	-0.24	6.47	89.66
11.0	69371443	0.983	0.001	0.000	0.960	0.007				3.02	-0.57	-0.25	6.28	89.65
11.3	70758872	0.983	0.001	0.001	0.961	0.007				2.92	-0.58	-0.25	6.10	89.64
11.5	72174049	0.983	0.001	0.001	0.961	0.007				2.82	-0.59	-0.26	5.91	89.63
11.7	73617530	0.983	0.001	0.001	0.962	0.007				2.72	-0.60	-0.26	5.73	89.62
12.0	75089881	0.983	0.001	0.001	0.962	0.007				2.63	-0.62	-0.27	5.55	89.61
12.2	76591678	0.984	0.001	0.001	0.963	0.007				2.53	-0.63	-0.27	5.37	89.60
12.4	78123512	0.984	0.001	0.001	0.963	0.007				2.43	-0.64	-0.28	5.19	89.58
12.7	79685982	0.984	0.001	0.001	0.964	0.008				2.34	-0.65	-0.29	5.02	89.57
12.9	81279702	0.984	0.001	0.001	0.964	0.008				2.25	-0.67	-0.29	4.85	89.56
13.2	82905296	0.984	0.001	0.001	0.965	0.008				2.15	-0.68	-0.30	4.68	89.55
13.5	84563402	0.985	0.001	0.001	0.965	0.008				2.06	-0.69	-0.30	4.51	89.54
13.7	86254670	0.985	0.001	0.001	0.966	0.008				1.97	-0.71	-0.31	4.34	89.53
14.0	87979763	0.985	0.001	0.001	0.966	0.008				1.88	-0.72	-0.32	4.17	89.51
14.3	89739358	0.985	0.001	0.001	0.967	0.009				1.79	-0.74	-0.32	4.01	89.50
14.6	91534146	0.985	0.001	0.001	0.967	0.009				1.70	-0.75	-0.33	3.85	89.49
14.9	93364828	0.986	0.001	0.001	0.968	0.009				1.62	-0.77	-0.34	3.68	89.48
15.2	95232125	0.986	0.001	0.001	0.968	0.009				1.53	-0.78	-0.34	3.52	89.46
15.5	97136768	0.986	0.001	0.001	0.969	0.009				1.44	-0.80	-0.35	3.36	89.45
15.8	99079503	0.987	0.001	0.001	0.969	0.010				1.35	-0.81	-0.36	3.21	89.44
16.1	101061093	0.987	0.001	0.001	0.970	0.010				1.27	-0.83	-0.36	3.05	89.42
16.4	103082315	0.987	0.001	0.001	0.970	0.010				1.18	-0.85	-0.37	2.89	89.41
16.7	105143961	0.987	0.001	0.001	0.971	0.010				1.10	-0.86	-0.38	2.74	89.39
17.1	107246840	0.988	0.002	0.001	0.972	0.010				1.01	-0.88	-0.39	2.59	89.38
17.4	109391777	0.988	0.002	0.001	0.972	0.011				0.93	-0.90	-0.39	2.43	89.37
17.8	111579613	0.988	0.002	0.001	0.973	0.011				0.85	-0.92	-0.40	2.28	89.35
18.1	113811205	0.989	0.002	0.001	0.973	0.011				0.76	-0.94	-0.41	2.13	89.34
18.5	116087429	0.989	0.002	0.001	0.974	0.011				0.68	-0.96	-0.42	1.98	89.32
18.8	118409178	0.990	0.002	0.001	0.975	0.011				0.60	-0.97	-0.43	1.83	89.30
19.2	120777361	0.990	0.002	0.001	0.975	0.012				0.51	-0.99	-0.44	1.68	89.29
19.6	123192908	0.990	0.002	0.002	0.976	0.012				0.43	-1.01	-0.44	1.53	89.27
20.0	125656767	0.991	0.002	0.002	0.977	0.012				0.35	-1.04	-0.45	1.38	89.25
20.4	128169902	0.991	0.002	0.002	0.978	0.012				0.26	-1.06	-0.46	1.23	89.24
20.8	130733300	0.992	0.002	0.002	0.978	0.013				0.18	-1.08	-0.47	1.09	89.22
21.2	133347966	0.992	0.002	0.002	0.979	0.013				0.10	-1.10	-0.48	0.94	89.20
21.6	136014925	0.993	0.002	0.002	0.980	0.013				0.02	-1.12	-0.49	0.79	89.18
22.1	138735224	0.993	0.003	0.002	0.981	0.013				-0.07	-1.15	-0.50	0.64	89.16
22.5	141509928	0.994	0.003	0.002	0.982	0.014				-0.15	-1.17	-0.51	0.50	89.15
23.0	144340127	0.994	0.003	0.002	0.983	0.014				-0.23	-1.19	-0.52	0.35	89.13
23.4	147226929	0.995	0.003	0.002	0.984	0.014				-0.31	-1.22	-0.53	0.20	89.11
23.9	150171468	0.996	0.003	0.002	0.985	0.015				-0.40	-1.24	-0.54	0.05	89.09
24.4	153174897	0.996	0.003	0.002	0.986	0.015				-0.48	-1.27	-0.55	-0.09	89.07
24.9	156238395	0.997	0.003	0.003	0.987	0.015				-0.57	-1.29	-0.57	-0.24	89.04
25.4	159363163	0.998	0.003	0.003	0.988	0.016				-0.65	-1.32	-0.58	-0.39	89.02
25.9	162550426	0.998	0.004	0.003	0.989	0.016				-0.74	-1.35	-0.59	-0.54	89.00
26.4	165801435	0.999	0.004	0.003	0.990	0.016				-0.82	-1.38	-0.60	-0.69	88.98
26.9	169117464	1.000	0.004	0.003	0.992	0.017				-0.91	-1.40	-0.61	-0.84	88.95
27.5	172499813	1.001	0.004	0.003	0.993	0.017				-0.99	-1.43	-0.63	-0.99	88.93
28.0	175949809	1.002	0.004	0.003	0.994	0.017				-1.08	-1.46	-0.64	-1.14	88.91
28.6	179468805	1.003	0.004	0.003	0.996	0.018				-1.17	-1.49	-0.65	-1.30	88.88
29.1	183058181	1.003	0.005	0.003	0.997	0.018				-1.25	-1.52	-0.66	-1.45	88.86
29.7	186719345	1.004	0.005	0.004	0.999	0.018				-1.34	-1.56	-0.68	-1.61	88.83
30.3	190453732	1.005	0.005	0.004	1.000	0.019				-1.43	-1.59	-0.69	-1.76	88.80
30.9	194262807	1.007	0.005	0.004	1.002	0.019				-1.52	-1.62	-0.71	-1.92	88.78
31.5	198148063	1.008	0.005	0.004	1.004	0.020				-1.61	-1.66	-0.72	-2.08	88.75
32.2	202111024	1.009	0.006	0.004	1.006	0.020				-1.71	-1.69	-0.74	-2.24	88.72
32.8	206153244	1.010	0.006	0.004	1.008	0.021				-1.80	-1.73	-0.75	-2.40	88.69
33.5	210276309	1.011	0.006	0.005	1.009	0.021				-1.89	-1.76	-0.77	-2.57	88.66
34.1	214481835	1.013	0.006	0.005	1.012	0.022				-1.99	-1.80	-0.78	-2.73	88.63
34.8	218771472	1.014	0.007	0.005	1.014	0.022				-2.08	-1.84	-0.80	-2.90	88.60
35.5	223146902	1.015	0.007	0.005	1.016	0.022				-2.18	-1.88	-0.82	-3.07	88.57
36.2	227609840	1.017	0.007	0.006	1.018	0.023				-2.28	-1.92	-0.83	-3.24	88.53
36.9	232162036	1.018	0.008	0.006	1.021	0.023				-2.38	-1.96	-0.85	-3.42	88.50
37.7	236805277	1.020	0.008	0.006	1.023	0.024				-2.48	-2.00	-0.87	-3.60	88.47
38.4	241541383	1.022	0.008	0.006	1.026	0.025				-2.59	-2.04	-0.89	-3.78	88.43
39.2	246372210	1.023	0.009	0.007	1.029	0.025				-2.69	-2.08	-0.90	-3.96	88.39
40.0	251299655	1.025	0.009	0.007	1.031	0.026				-2.80	-2.13	-0.92	-4.15	88.36
40.8	256325648	1.027	0.009	0.007	1.034	0.026				-2.90	-2.18	-0.94	-4.34	88.32
41.6	261452161	1.029	0.010	0.007	1.038	0.027				-3.01	-2.22	-0.96	-4.53	88.28
42.4	266681204	1.031	0.010	0.008	1.041	0.028				-3.13	-2.27	-0.98	-4.72	88.24
43.3	272014828	1.033	0.011	0.008	1.044	0.028				-3.24	-2.32	-1.00	-4.92	88.19
44.2	277455125	1.036	0.011	0.008	1.048	0.029				-3.35	-2.37	-1.03	-5.13	88.15
45.0	283004227	1.038	0.012	0.009	1.052	0.030				-3.47	-2.42	-1.05	-5.34	88.10
45.9	288664312	1.041	0.012	0.009	1.056	0.030				-3.59	-2.48	-1.07	-5.55	88.06
46.9	294437598	1.043	0.013	0.010	1.060	0.031				-3.71	-2.53	-1.09	-5.77	88.01
47.8	300326350	1.046	0.014	0.010	1.064	0.032				-3.84	-2.59	-1.12	-5.99	87.96
48.8	306332877	1.049	0.014	0.011	1.069	0.033				-3.97	-2.65	-1.14	-6.22	87.91
49.7	312459534	1.052	0.015	0.011	1.073	0.033				-4.10	-2.71	-1.17	-6.45	87.86
50.7	318708725	1.055	0.016	0.011	1.078	0.034				-4.23	-2.77	-1.19	-6.69	87.80
51.7	325082899	1.058	0.016	0.012	1.084	0.035				-4.37	-2.83	-1.22	-6.94	87.75
52.8	331584557	1.061	0.017	0.013	1.089	0.036				-4.50	-2.89	-1.24	-7.19	87.69
53.8	338216249	1.065	0.018	0.013	1.095	0.037				-4.65	-2.96	-1.27	-7.45	87.63
54.9	344980574	1.069	0.019	0.014	1.101	0.038				-4.79	-3.03	-1.30	-7.71	87.57
56.0	351880185	1.073	0.020	0.014	1.107	0.039				-4.94	-3.10	-1.33	-7.99	87.50
57.1	358917789	1.077	0.021	0.015	1.114	0.040				-5.10	-3.17	-1.36	-8.27	87.44
58.3	366096144	1.081	0.022	0.016	1.121	0.041				-5.25	-3.25	-1.39	-8.56	87.37
59.4	373418067	1.086	0.023	0.017	1.129	0.042				-5.41	-3.32	-1.42	-8.87	87.29
60.6	380886429	1.090	0.024	0.017	1.136	0.043				-5.58	-3.40	-1.45	-9.18	87.22
61.8	388504157	1.095	0.025	0.018	1.145	0.044				-5.75	-3.48	-1.49	-9.50	87.14
63.1	396274240	1.100	0.026	0.019	1.153	0.046				-5.93	-3.57	-1.52	-9.83	87.06
64.3	404199725	1.106	0.028	0.020	1.162	0.047				-6.11	-3.66	-1.56	-10.18	86.97
65.6	412283720	1.112	0.029	0.021	1.172	0.048				-6.30	-3.75	-1.59	-10.54	86.89
66.9	420529394	1.118	0.031	0.022	1.182	0.050				-6.49	-3.84	-1.63	-10.91	86.79
68.3	428939982	1.124	0.033	0.023	1.193	0.051				-6.69	-3.94	-1.67	-11.30	86.70
69.6	437518782	1.131	0.034	0.024	1.204	0.053				-6.89	-4.04	-1.71	-11.71	86.60
71.0	446269157	1.138	0.036	0.026	1.216	0.054				-7.10	-4.14	-1.75	-12.13	86.49
72.4	455194540	1.145	0.038	0.027	1.229	0.056				-7.32	-4.25	-1.79	-12.57	86.38
73.9	464298431	1.153	0.040	0.028	1.242	0.058				-7.55	-4.36	-1.84	-13.04	86.26
75.4	473584400	1.161	0.043	0.030	1.257	0.060				-7.79	-4.48	-1.88	-13.52	86.14
76.9	483056088	1.169	0.045	0.032	1.272	0.062				-8.03	-4.60	-1.93	-14.03	86.01
78.4	492717210	1.179	0.048	0.033	1.288	0.064				-8.28	-4.73	-1.98	-14.56	85.87
80.0	502571554	1.188	0.051	0.035	1.305	0.066				-8.55	-4.86	-2.03	-15.13	85.72
81.6	512622985	1.198	0.054	0.037	1.323	0.068				-8.82	-5.00	-2.08	-15.72	85.57
83.2	522875445	1.209	0.058	0.039	1.342	0.071				-9.11	-5.15	-2.14	-16.34	85.40
84.9	533332954	1.220	0.061	0.042	1.363	0.074				-9.41	-5.30	-2.20	-17.01	85.23
86.6	543999613	1.232	0.066	0.044	1.385	0.077				-9.72	-5.46	-2.26	-17.71	85.04
88.3	554879605	1.245	0.070	0.047	1.409	0.080				-10.05	-5.63	-2.32	-18.45	84.84
90.1	565977197	1.259	0.075	0.050	1.434	0.083				-10.39	-5.81	-2.38	-19.25	84.63
91.9	577296741	1.273	0.080	0.053	1.461	0.086				-10.75	-6.00	-2.45	-20.10	84.40
93.7	588842676	1.288	0.086	0.056	1.489	0.090				-11.12	-6.20	-2.53	-21.00	84.15
95.6	600619529	1.304	0.093	0.060	1.520	0.094				-11.52	-6.41	-2.60	-21.97	83.89
97.5	612631920	1.322	0.100	0.064	1.553	0.099				-11.93	-6.64	-2.68	-23.02	83.59
99.5	624884558	1.340	0.108	0.068	1.589	0.104				-12.37	-6.88	-2.77	-24.14	83.28
101.4	637382249	1.360	0.117	0.073	1.627	0.109				-12.84	-7.14	-2.86	-25.35	82.93
103.5	650129894	1.381	0.128	0.078	1.668	0.115				-13.33	-7.42	-2.95	-26.67	82.55
105.5	663132492	1.403	0.139	0.084	1.712	0.121				-13.86	-7.73	-3.06	-28.09	82.13
107.7	676395142	1.427	0.153	0.091	1.760	0.128				-14.41	-8.07	-3.17	-29.65	81.66
109.8	689923045	1.453	0.168	0.098	1.811	0.136				-15.00	-8.44	-3.29	-31.35	81.14
112.0	703721506	1.481	0.186	0.107	1.865	0.145				-15.64	-8.85	-3.42	-33.21	80.55
114.2	717795936	1.511	0.207	0.116	1.923	0.155				-16.31	-9.31	-3.56	-35.25	79.89
116.5	732151855	1.543	0.232	0.127	1.985	0.167				-17.04	-9.84	-3.71	-37.50	79.13
118.9	746794892	1.578	0.261	0.139	2.051	0.180				-17.82	-10.44	-3.88	-39.98	78.25
119.5	750528866	1.587	0.270	0.142	2.068	0.184				-18.03	-10.61	-3.93	-40.65	78.01
120.0	754281511	1.596	0.279	0.146	2.085	0.188				-18.24	-10.78	-3.97	-41.33	77.75
120.6	758052918	1.606	0.288	0.150	2.102	0.192				-18.45	-10.96	-4.02	-42.03	77.49
121.3	761843183	1.615	0.298	0.153	2.120	0.196				-18.67	-11.15	-4.07	-42.75	77.22
121.9	765652399	1.625	0.309	0.157	2.138	0.200				-18.89	-11.35	-4.12	-43.48	76.93
122.5	769480661	1.636	0.320	0.161	2.156	0.204				-19.12	-11.55	-4.18	-44.24	76.63
123.1	773328064	1.646	0.332	0.166	2.174	0.209				-19.35	-11.77	-4.23	-45.01	76.32
123.7	777194704	1.657	0.344	0.170	2.192	0.214				-19.59	-12.00	-4.29	-45.81	76.00
124.3	781080678	1.667	0.357	0.175	2.210	0.219				-19.83	-12.24	-4.35	-46.62	75.66
124.9	784986081	1.678	0.372	0.180	2.228	0.224				-20.08	-12.49	-4.41	-47.46	75.30
125.6	788911012	1.690	0.387	0.185	2.246	0.230				-20.33	-12.76	-4.47	-48.32	74.92
126.2	792855567	1.701	0.403	0.190	2.265	0.236				-20.59	-13.04	-4.54	-49.20	74.53
126.8	796819844	1.713	0.420	0.195	2.283	0.242				-20.86	-13.34	-4.61	-50.10	74.12
127.5	800803944	1.725	0.439	0.201	2.301	0.248				-21.13	-13.66	-4.68	-51.03	73.68
128.1	804807963	1.737	0.459	0.207	2.319	0.255				-21.40	-14.00	-4.75	-51.97	73.22
128.7	808832003	1.750	0.481	0.214	2.337	0.262				-21.68	-14.36	-4.83	-52.95	72.74
129.4	812876163	1.763	0.504	0.220	2.355	0.270				-21.97	-14.75	-4.90	-53.94	72.23
130.0	816940544	1.776	0.529	0.228	2.372	0.278				-22.27	-15.17	-4.99	-54.96	71.69
130.7	821025247	1.789	0.557	0.235	2.389	0.286				-22.57	-15.62	-5.07	-56.01	71.12
131.3	825130373	1.803	0.587	0.243	2.406	0.295				-22.88	-16.11	-5.17	-57.08	70.51
132.0	829256025	1.817	0.620	0.251	2.423	0.304				-23.20	-16.64	-5.26	-58.17	69.86
132.6	833402305	1.831	0.657	0.260	2.438	0.314				-23.52	-17.22	-5.36	-59.29	69.17
133.3	837569316	1.845	0.697	0.269	2.454	0.325				-23.85	-17.86	-5.47	-60.44	68.44
134.0	841757163	1.860	0.742	0.279	2.469	0.336				-24.19	-18.56	-5.58	-61.61	67.66
134.6	845965949	1.876	0.792	0.290	2.483	0.348				-24.54	-19.35	-5.69	-62.80	66.82
135.3	850195779	1.891	0.849	0.301	2.496	0.361				-24.90	-20.22	-5.82	-64.02	65.92
136.0	854446757	1.907	0.912	0.313	2.508	0.374				-25.26	-21.21	-5.95	-65.27	64.95
136.7	858718991	1.924	0.985	0.326	2.520	0.389				-25.64	-22.32	-6.09	-66.54	63.91
137.4	863012586	1.940	1.068	0.340	2.530	0.404				-26.02	-23.60	-6.24	-67.83	62.78
138.0	867327649	1.957	1.165	0.355	2.540	0.421				-26.42	-25.08	-6.39	-69.14	61.56
138.7	871664287	1.975	1.277	0.371	2.548	0.439				-26.82	-26.82	-6.56	-70.48	60.24
139.4	876022609	1.993	1.410	0.388	2.555	0.458				-27.24	-28.87	-6.74	-71.84	58.80
140.1	880402722	2.011	1.568	0.407	2.561	0.478				-27.67	-31.34	-6.94	-73.21	57.24
140.8	884804735	2.030	1.758	0.427	2.566	0.500				-28.11	-34.35	-7.15	-74.61	55.53
141.5	889228759	2.049	1.990	0.449	2.569	0.524				-28.56	-38.10	-7.38	-76.03	53.66
142.2	893674903	2.068	2.272	0.474	2.570	0.549				-29.02	-42.85	-7.63	-77.46	51.60
142.9	898143277	2.088	2.614	0.500	2.570	0.576				-29.50	-48.97	-7.91	-78.90	49.35
143.7	902633994	2.109	3.016	0.530	2.569	0.604				-29.99	-56.96	-8.21	-80.36	46.86
144.4	907147164	2.130	3.454	0.563	2.565	0.635				-30.50	-67.34	-8.54	-81.83	44.13
145.1	911682900	2.151	3.844	0.600	2.561	0.666				-31.02	-80.37	-8.91	-83.31	41.11
145.8	916241314	2.173	4.051	0.641	2.554	0.699				-31.55	84.58	-9.32	-84.80	37.80
146.6	920822521	2.196	3.979	0.688	2.546	0.732				-32.10	69.26	-9.79	-86.29	34.15
147.3	925426633	2.219	3.678	0.742	2.536	0.764				-32.67	55.61	-10.32	-87.79	30.17
148.0	930053766	2.242	3.284	0.803	2.524	0.796				-33.25	44.56	-10.93	-89.29	25.83
148.8	934704035	2.266	2.897	0.875	2.511	0.824				-33.85	36.01	-11.64	89.21	21.15
149.5	939377556	2.290	2.558	0.960	2.496	0.849				-34.47	29.47	-12.47	87.71	16.16
150.3	944074443	2.315	2.275	1.061	2.479	0.868				-35.11	24.40	-13.46	86.21	10.91
151.0	948794815	2.341	2.041	1.183	2.461	0.880				-35.77	20.42	-14.67	84.72	5.46
151.8	953538790	2.367	1.848	1.335	2.442	0.884				-36.45	17.23	-16.16	83.24	-0.09
152.5	958306484	2.394	1.687	1.527	2.421	0.880				-37.15	14.62	-18.06	81.77	-5.63
153.3	963098016	2.421	1.552	1.777	2.399	0.867				-37.88	12.45	-20.55	80.31	-11.08
154.0	967913506	2.448	1.438	2.113	2.375	0.848				-38.62	10.63	-23.94	78.86	-16.32
154.8	972753074	2.476	1.340	2.583	2.350	0.824				-39.39	9.07	-28.80	77.43	-21.30
155.6	977616839	2.505	1.256	3.264	2.324	0.795				-40.19	7.72	-36.23	76.01	-25.97
156.4	982504923	2.534	1.182	4.255	2.297	0.763				-41.01	6.54	-48.34	74.61	-30.29
157.2	987417448	2.564	1.118	5.475	2.270	0.730				-41.86	5.50	-68.66	73.23	-34.27
157.9	992354535	2.594	1.061	6.019	2.241	0.697				-42.74	4.57	82.34	71.87	-37.90
158.7	997316308	2.624	1.011	5.174	2.212	0.665				-43.64	3.73	55.45	70.53	-41.21
159.5	1002302889	2.655	0.966	4.015	2.182	0.634				-44.58	2.98	38.12	69.22	-44.21
160.3	1007314404	2.686	0.925	3.146	2.151	0.603				-45.55	2.29	27.81	67.92	-46.94
161.1	1012350976	2.718	0.889	2.550	2.120	0.575				-46.54	1.65	21.36	66.65	-49.42
161.9	1017412731	2.749	0.856	2.133	2.089	0.548				-47.57	1.07	17.04	65.40	-51.67
162.7	1022499794	2.781	0.826	1.831	2.057	0.523				-48.64	0.53	13.98	64.18	-53.71
163.5	1027612293	2.813	0.798	1.603	2.025	0.500				-49.74	0.03	11.70	62.98	-55.58
164.4	1032750355	2.845	0.773	1.426	1.993	0.478				-50.88	-0.45	9.94	61.81	-57.28
165.2	1037914106	2.877	0.750	1.286	1.961	0.457				-52.05	-0.89	8.55	60.66	-58.84
166.0	1043103677	2.908	0.729	1.171	1.929	0.438				-53.26	-1.31	7.42	59.54	-60.28
166.8	1048319195	2.940	0.710	1.076	1.897	0.420				-54.51	-1.71	6.48	58.44	-61.59
167.7	1053560791	2.971	0.692	0.996	1.865	0.404				-55.79	-2.09	5.68	57.36	-62.81
168.5	1058828595	3.001	0.675	0.928	1.834	0.388				-57.12	-2.45	5.00	56.32	-63.93
169.4	1064122738	3.031	0.659	0.870	1.802	0.374				-58.48	-2.79	4.42	55.29	-64.97
170.2	1069443352	3.060	0.645	0.819	1.771	0.360				-59.89	-3.13	3.90	54.29	-65.94
171.1	1074790569	3.088	0.632	0.774	1.740	0.348				-61.34	-3.45	3.45	53.32	-66.84
171.9	1080164522	3.114	0.619	0.734	1.709	0.336				-62.82	-3.76	3.04	52.37	-67.67
172.8	1085565344	3.139	0.607	0.699	1.679	0.325				-64.35	-4.06	2.68	51.44	-68.46
173.6	1090993171	3.163	0.597	0.667	1.649	0.314				-65.91	-4.35	2.35	50.53	-69.19
174.5	1096448137	3.185	0.586	0.638	1.620	0.304				-67.51	-4.64	2.05	49.65	-69.87
175.4	1101930377	3.205	0.577	0.613	1.590	0.295				-69.15	-4.91	1.78	48.79	-70.52
176.3	1107440029	3.223	0.568	0.589	1.562	0.286				-70.82	-5.19	1.53	47.95	-71.12
177.1	1112977229	3.238	0.560	0.568	1.534	0.278				-72.53	-5.45	1.30	47.14	-71.70
178.0	1118542116	3.251	0.552	0.548	1.506	0.270				-74.27	-5.72	1.08	46.34	-72.24
181.6	1140912958	3.272	0.525	0.484	1.400	0.242				-81.44	-6.73	0.36	43.37	-74.11
185.2	1163731217	3.241	0.505	0.436	1.302	0.219				-88.88	-7.70	-0.21	40.68	-75.63
188.9	1187005841	3.157	0.490	0.399	1.212	0.200				83.66	-8.68	-0.68	38.26	-76.90
192.7	1210745958	3.026	0.479	0.370	1.129	0.184				76.41	-9.67	-1.09	36.06	-77.97
196.6	1234960877	2.858	0.471	0.346	1.052	0.170				69.57	-10.69	-1.44	34.06	-78.89
200.5	1259660095	2.670	0.467	0.326	0.982	0.158				63.29	-11.77	-1.76	32.24	-79.68
204.5	1284853297	2.474	0.466	0.310	0.918	0.148				57.63	-12.92	-2.06	30.58	-80.37
208.6	1310550363	2.280	0.468	0.297	0.859	0.139				52.59	-14.16	-2.34	29.06	-80.97
212.8	1336761370	2.095	0.472	0.286	0.804	0.130				48.15	-15.54	-2.60	27.66	-81.51
217.0	1363496597	1.923	0.480	0.276	0.754	0.123				44.23	-17.06	-2.86	26.38	-81.99
221.3	1390766529	1.765	0.492	0.268	0.708	0.116				40.79	-18.78	-3.11	25.19	-82.43
225.8	1418581860	1.621	0.507	0.261	0.665	0.110				37.77	-20.75	-3.36	24.10	-82.82
230.3	1446953497	1.492	0.526	0.256	0.626	0.105				35.09	-23.03	-3.61	23.08	-83.17
234.9	1475892567	1.375	0.550	0.251	0.590	0.100				32.73	-25.70	-3.86	22.13	-83.50
239.6	1505410418	1.269	0.579	0.247	0.556	0.095				30.62	-28.88	-4.12	21.25	-83.80
244.4	1535518627	1.174	0.615	0.244	0.525	0.091				28.74	-32.71	-4.38	20.42	-84.07
249.3	1566228999	1.089	0.657	0.242	0.495	0.087				27.05	-37.40	-4.65	19.65	-84.32
254.3	1597553579	1.011	0.705	0.241	0.468	0.084				25.53	-43.20	-4.93	18.93	-84.56
259.3	1629504651	0.941	0.757	0.240	0.443	0.080				24.15	-50.40	-5.23	18.25	-84.77
264.5	1662094744	0.877	0.807	0.239	0.419	0.077				22.90	-59.29	-5.54	17.61	-84.98
269.8	1695336639	0.819	0.845	0.240	0.397	0.074				21.76	-69.95	-5.87	17.00	-85.17
275.2	1729243372	0.766	0.854	0.241	0.377	0.072				20.72	-82.07	-6.21	16.43	-85.34
280.7	1763828239	0.718	0.825	0.243	0.357	0.069				19.76	85.22	-6.59	15.89	-85.51
286.3	1799104804	0.673	0.761	0.245	0.339	0.067				18.88	73.10	-6.99	15.38	-85.67
292.1	1835086900	0.632	0.678	0.248	0.322	0.064				18.06	62.43	-7.42	14.89	-85.82
297.9	1871788638	0.594	0.593	0.253	0.306	0.062				17.31	53.56	-7.89	14.43	-85.96
303.9	1909224411	0.560	0.515	0.257	0.291	0.060				16.60	46.37	-8.40	13.99	-86.09
309.9	1947408899	0.528	0.446	0.263	0.277	0.058				15.95	40.60	-8.97	13.57	-86.21
316.1	1986357077	0.498	0.389	0.270	0.264	0.057				15.34	35.94	-9.60	13.17	-86.33
322.5	2026084218	0.470	0.341	0.279	0.251	0.055				14.76	32.15	-10.30	12.79	-86.44
328.9	2066605903	0.444	0.301	0.289	0.239	0.053				14.22	29.02	-11.09	12.42	-86.55
335.5	2107938021	0.420	0.267	0.300	0.228	0.052				13.72	26.41	-11.99	12.07	-86.65
342.2	2150096781	0.398	0.238	0.314	0.217	0.050				13.24	24.21	-13.02	11.74	-86.75
349.0	2193098717	0.377	0.214	0.330	0.207	0.049				12.79	22.33	-14.21	11.42	-86.85
356.0	2236960691	0.358	0.193	0.350	0.198	0.047				12.37	20.70	-15.61	11.11	-86.93
363.1	2281699905	0.339	0.175	0.373	0.189	0.046				11.97	19.29	-17.28	10.81	-87.02
370.4	2327333903	0.322	0.159	0.402	0.180	0.045				11.58	18.05	-19.30	10.53	-87.10
377.8	2373880581	0.306	0.146	0.437	0.172	0.043				11.22	16.95	-21.80	10.25	-87.18
385.4	2421358193	0.291	0.134	0.481	0.164	0.042				10.88	15.97	-24.94	9.99	-87.26
393.1	2469785357	0.277	0.123	0.536	0.157	0.041				10.55	15.09	-29.00	9.73	-87.33
400.9	2519181064	0.263	0.113	0.605	0.150	0.040				10.24	14.30	-34.38	9.49	-87.40
409.0	2569564685	0.251	0.105	0.692	0.143	0.039				9.94	13.58	-41.72	9.25	-87.47
417.1	2620955979	0.239	0.097	0.795	0.137	0.038				9.65	12.93	-51.88	9.02	-87.53
425.5	2673375098	0.227	0.090	0.895	0.131	0.037				9.38	12.33	-65.77	8.80	-87.59
434.0	2726842600	0.217	0.084	0.947	0.125	0.036				9.12	11.78	-83.30	8.58	-87.65
442.7	2781379452	0.207	0.078	0.906	0.120	0.035				8.87	11.27	77.96	8.38	-87.71
451.5	2837007041	0.197	0.073	0.792	0.115	0.034				8.63	10.81	61.52	8.17	-87.77
460.6	2893747182	0.188	0.069	0.661	0.110	0.034				8.39	10.37	48.97	7.98	-87.82
469.8	2951622126	0.180	0.064	0.547	0.105	0.033				8.17	9.97	39.87	7.79	-87.88
479.2	3010654568	0.171	0.060	0.456	0.101	0.032				7.96	9.59	33.28	7.61	-87.93
488.7	3070867660	0.164	0.057	0.385	0.097	0.031				7.75	9.23	28.41	7.43	-87.98
498.5	3132285013	0.156	0.053	0.329	0.092	0.030				7.55	8.90	24.70	7.26	-88.02
508.5	3194930713	0.149	0.050	0.285	0.089	0.030				7.36	8.59	21.81	7.09	-88.07
518.7	3258829327	0.143	0.047	0.250	0.085	0.029				7.18	8.30	19.50	6.93	-88.12
529.0	3324005914	0.137	0.045	0.221	0.081	0.028				7.00	8.02	17.62	6.77	-88.16
539.6	3390486032	0.131	0.042	0.197	0.078	0.028				6.83	7.76	16.05	6.62	-88.20
550.4	3458295753	0.125	0.040	0.177	0.075	0.027				6.66	7.51	14.74	6.47	-88.24
561.4	3527461668	0.120	0.038	0.159	0.072	0.026				6.50	7.27	13.61	6.33	-88.28
572.6	3598010901	0.114	0.036	0.145	0.069	0.026				6.34	7.05	12.65	6.19	-88.32
584.1	3669971119	0.109	0.034	0.132	0.066	0.025				6.19	6.84	11.80	6.05	-88.36
595.8	3743370542	0.105	0.032	0.121	0.063	0.025				6.04	6.63	11.06	5.92	-88.40
607.7	3818237952	0.100	0.031	0.111	0.060	0.024				5.90	6.44	10.40	5.79	-88.43
619.8	3894602711	0.096	0.029	0.102	0.058	0.024				5.77	6.25	9.81	5.66	-88.47
632.2	3972494766	0.092	0.028	0.095	0.056	0.023				5.63	6.08	9.28	5.54	-88.50
644.9	4051944661	0.088	0.026	0.088	0.053	0.023				5.50	5.91	8.80	5.42	-88.53
657.8	4132983554	0.084	0.025	0.081	0.051	0.022				5.38	5.74	8.37	5.30	-88.56
670.9	4215643225	0.081	0.024	0.076	0.049	0.022				5.25	5.59	7.97	5.19	-88.60
684.4	4299956090	0.078	0.023	0.071	0.047	0.021				5.14	5.44	7.61	5.08	-88.63
698.0	4385955212	0.074	0.021	0.066	0.045	0.021				5.02	5.29	7.28	4.97	-88.66
712.0	4473674316	0.071	0.020	0.062	0.043	0.020				4.91	5.16	6.97	4.86	-88.68
726.2	4563147802	0.068	0.019	0.058	0.042	0.020				4.80	5.02	6.68	4.76	-88.71
740.8	4654410758	0.065	0.019	0.055	0.040	0.019				4.69	4.89	6.42	4.66	-88.74
755.6	4747498973	0.063	0.018	0.051	0.038	0.019				4.59	4.77	6.17	4.56	-88.77
770.7	4842448953	0.060	0.017	0.048	0.037	0.019				4.49	4.65	5.94	4.47	-88.79
786.1	4939297932	0.058	0.016	0.046	0.035	0.018				4.39	4.54	5.72	4.37	-88.82
801.8	5038083891	0.055	0.015	0.043	0.034	0.018				4.30	4.42	5.52	4.28	-88.84
817.9	5138845568	0.053	0.015	0.041	0.033	0.017				4.20	4.32	5.33	4.19	-88.87
834.2	5241622480	0.051	0.014	0.038	0.031	0.017				4.11	4.21	5.15	4.10	-88.89
850.9	5346454929	0.049	0.013	0.036	0.030	0.017				4.02	4.11	4.98	4.02	-88.92
867.9	5453384028	0.047	0.013	0.034	0.029	0.016				3.94	4.01	4.82	3.94	-88.94
885.3	5562451708	0.045	0.012	0.033	0.028	0.016				3.85	3.92	4.66	3.86	-88.96
903.0	5673700743	0.043	0.012	0.031	0.027	0.016				3.77	3.83	4.52	3.78	-88.98
921.1	5787174757	0.041	0.011	0.029	0.025	0.015				3.69	3.74	4.38	3.70	-89.00
939.5	5902918253	0.040	0.011	0.028	0.024	0.015				3.61	3.65	4.25	3.62	-89.02
958.3	6020976618	0.038	0.010	0.026	0.024	0.015				3.54	3.57	4.12	3.55	-89.04
977.4	6141396150	0.037	0.010	0.025	0.023	0.014				3.46	3.49	4.00	3.48	-89.06
997.0	6264224073	0.035	0.009	0.024	0.022	0.014				3.39	3.41	3.89	3.40	-89.08
1016.9	6389508554	0.034	0.009	0.023	0.021	0.014				3.32	3.33	3.78	3.33	-89.10

## Light
Photoresistance	100	Ohm

EZ290_B_1				EZ500_Gr_1				EZ290_Gr_1				EZ290_Gr_2				EZ290_Bl_2				EZ290_Bl_3					LED			Fiber
Wavelength	470	nm		Wavelength	527	nm		Wavelength	527	nm		Wavelength	527	nm		Wavelength	470	nm		Wavelength	470	nm		Current	30	mA		Current	30	mA	Current		Fiber	Dia		Size	Color	Power	Power	Capture	Theory
Response	0.2	A/W		Response	0.25	A/W		Response	0.25	A/W		Response	0.25	A/W		Response	0.2	A/W		Response	0.2	A/W		Type	C460EZ500			Type	C470EZ290		(mA)		Face	(um)	NA	(um)	(nm)	(mW)	(mW)	(%)	(%)	Comments
							Wavelength	460	nm		Wavelength	470	nm	0.49		1A	300	0.4138029443	480	460	26	2.5	9.4	5.3		1.8
Diode	Photovoltage	Power		Diode	Photovoltage	Power		Diode	Photovoltage	Power		Diode	Photovoltage	Power		Diode	Photovoltage	Power		Diode	Photovoltage	Power			Photocurrent	Power			Photocurrent	Power	0.49		1B	300	0.4138029443	480	460	26	2.5	9.4	5.3		1.8
(mA)	(V)	(mW)		(mA)	(V)	(mW)		(mA)	(V)	(mW)		(mA)	(V)	(mW)		(mA)	(V)	(mW)		(mA)	(V)	(mW)		LED	(mA)	(mW)		LED	(mA)	(mW)	0.41		2A	300	0.4138029443	480	460	26	2.1	7.9	5.3	Defects	1.5
0.0	0.00	0.00		0.0	0.00	0.00		0.0	0.00	0.00		0.0	0.00	0.00		0.0	0.00	0.00		0.0	0.00	0.00		1	5.17	25.85		1	3.28	16.40	0.42		2B	300	0.4138029443	480	460	26	2.1	8.1	5.3		1.5
2.4	0.04	2.00		7.8	0.12	4.68		4.6	0.05	1.92		4.7	0.05	1.84		2.3	0.03	1.70		6.0	0.09	4.45		2	5.03	25.15		2	3.59	17.95	0.50		3A	300	0.4138029443	480	460	26	2.5	9.6	5.3		1.8
6.1	0.10	4.80		13.0	0.19	7.40		11.0	0.12	4.68		7.7	0.08	3.00		6.2	0.09	4.65		11.3	0.16	8.00		3	5.02	25.10		3	3.60	18.00	0.51		3B	300	0.4138029443	480	460	26	2.6	9.8	5.3		1.9
10.9	0.15	7.70		16.0	0.22	8.76		15.8	0.15	6.08		11.6	0.11	4.20		8.0	0.12	5.90		14.6	0.20	10.05		4	5.12	25.60		4	3.64	18.20	0.32		1A	400	0.25	480	460	26	1.6	6.2	3.4		1.8
13.9	0.19	9.40		23.0	0.29	11.72		21.1	0.19	7.48		18.2	0.15	5.88		11.1	0.16	8.00		17.2	0.23	11.55		5	5.20	26.00		5	3.68	18.40	0.31		1B	400	0.2544932993	480	460	26	1.6	6.0	3.5		1.7
20.2	0.26	13.10		30.2	0.37	14.68		23.3	0.20	8.00		21.6	0.17	6.72		13.4	0.19	9.40		19.8	0.26	13.05			Average	25.54			Average	17.79	0.26		2A	400	0.2544932993	480	460	26	1.3	5.0	3.5		1.4
23.9	0.29	14.65		36.6	0.42	16.60		28.6	0.23	9.28		24.0	0.18	7.16		15.1	0.21	10.35		22.1	0.29	14.25					0.27		2B	400	0.2544932993	480	460	26	1.4	5.2	3.5	Not flat	1.5
27.8	0.33	16.55		41.5	0.45	18.12		32.2	0.25	9.96		25.9	0.19	7.56		19.0	0.25	12.60		25.2	0.32	15.85		Current	30	mA		Current	30	mA	0.30		3A	400	0.2544932993	480	460	26	1.5	5.8	3.5	Defects	1.6
33.5	0.38	18.90		47.0	0.49	19.72		35.5	0.266	10.64		28.2	0.20	8.00		23.2	0.30	14.75		28.0	0.33	16.55		Type	C527EZ500			Type	C527EZ290		0.30		3B	400	0.2544932993	480	460	26	1.5	5.8	3.5		1.6
37.9	0.41	20.50					40.2	0.288	11.52		30.5	0.212	8.48		25.9	0.321	16.05		30.5	0.365	18.25		Wavelength	527	nm		Wavelength	527	nm	0.26		1A	400	0.242535625	480	460	26	1.3	5.0	3.2		1.6
41.5	0.44	21.80					46.6	0.315	12.60		35.2	0.233	9.32		28.3	0.341	17.05		33.8	0.395	19.75			Photocurrent	Power			Photocurrent	Power	0.26		1B	400	0.242535625	480	460	26	1.3	5.0	3.2		1.6
45.8	0.46	22.85					40.9	0.257	10.28		30.0	0.355	17.75		38.2	0.431	21.55		LED	(mA)	(mW)		LED	(mA)	(mW)	0.25		2A	400	0.242535625	480	460	26	1.3	4.8	3.2		1.5
							46.6	0.279	11.16		35.0	0.393	19.65		42.0	0.46	23.00		1	3.48	13.92		1	2.32	9.28	0.25		2B	400	0.242535625	480	460	26	1.3	4.8	3.2		1.5
Current (mA) / Light Power (mW)							40.0	0.434	21.70				0.00		2	3.30	13.20		2	2.01	8.04	0.26		3A	400	0.242535625	480	460	26	1.3	5.0	3.2		1.6
							3	3.34	13.36			Average	8.66	0.25		3B	400	0.242535625	480	460	26	1.3	4.8	3.2		1.5
							4	3.38	13.52					0.53		1A	400	0.3713906764	480	460	26	2.7	10.2	7.5		1.4
							5	3.40	13.60					0.55		1B	400	0.3713906764	480	460	26	2.8	10.6	7.5		1.4
							6	3.42	13.68					0.56		2A	400	0.3713906764	480	460	26	2.8	10.8	7.5		1.4
							Average	13.55					0.57		2B	400	0.3713906764	480	460	26	2.9	11.0	7.5		1.5
							0.55		3A	400	0.3713906764	480	460	26	2.8	10.6	7.5		1.4
							0.56		3B	400	0.3713906764	480	460	26	2.8	10.8	7.5		1.4
							0.58		1A	300	0.4138029443	290	470	17	2.9	17.1	13.7		1.2
							0.58		1B	300	0.4138029443	290	470	17	2.9	17.1	13.7		1.2
							0.49		2A	300	0.4138029443	290	470	17	2.5	14.4	13.7	Defects	1.1
							0.47		2B	300	0.4138029443	290	470	17	2.4	13.8	13.7		1.0
							0.57		3A	300	0.4138029443	290	470	17	2.9	16.8	13.7		1.2
							0.59		3B	300	0.4138029443	290	470	17	3.0	17.4	13.7		1.3
							0.51		1A	400	0.3713906764	290	470	17	2.6	15.0	13.8		1.1
							0.52		1B	400	0.3713906764	290	470	17	2.6	15.3	13.8		1.1
							0.50		2A	400	0.3713906764	290	470	17	2.5	14.7	13.8		1.1
							0.52		2B	400	0.3713906764	290	470	17	2.6	15.3	13.8		1.1
							0.51		3A	400	0.3713906764	290	470	17	2.6	15.0	13.8		1.1
							0.52		3B	400	0.3713906764	290	470	17	2.6	15.3	13.8		1.1

A				B				C
Wavelength	460	nm		Wavelength	460	nm		Wavelength	460	nm
Response	0.182	A/W		Response	0.182	A/W		Response	0.182	A/W

Diode	Photocurrent	Power		Diode	Photocurrent	Power		Diode	Photocurrent	Power
(mA)	(mA)	(mW)		(mA)	(mA)	(mW)		(mA)	(mA)	(mW)
0.0	0.00	0.00		0.0	0.00	0.00		0.0	0.00	0.00
1.1	0.17	0.93		1.5	0.18	0.99		0.7	0.09	0.49
9.4	1.75	9.62		7.2	1.35	7.42		4.5	0.77	4.23
16.5	3.00	16.48		10.3	1.87	10.27		11.1	2.01	11.04
19.5	3.49	19.18		14.4	2.61	14.34		17.4	2.98	16.37
23.0	4.05	22.25		18.5	3.30	18.13		22.6	3.86	21.21
29.3	4.99	27.42		22.2	3.89	21.37		25.8	4.35	23.90
33.0	5.54	30.44		26.4	4.52	24.84		32.0	5.22	28.68
42.4	6.81	37.42		29.6	5.00	27.47		37.9	6.01	33.02
49.0	7.58	41.65		33.7	5.57	30.60		42.8	6.63	36.43
				37.3	6.05	33.24		49.7	7.47	41.04
				41.1	6.52	35.82
				48.8	7.44	40.88

Current (mA) / Light Power (mW)

## Resonance
	F_tank	C_tank			No6 Encapsulated
Circuit	(MHz)	(pF)	Environment		Test	T_low (ns)	T_hi (ns)	f_hi (MHz)	f_lo (MHz)	F_mid (MHz)
1	150.0	8.0	bare (frequency is approx)		1	6.66	7.06	150.1	141.6	145.8
1	146.0	8.0	liquid epoxy (frequency is approx)		2	6.54	7.05	152.8	141.9	147.4
1	150.8	8.0	cured epoxy		3	6.80	6.99	147.1	143.1	145.1
1	151.4	8.0	one week later		4	6.52	7.23	153.4	138.3	145.8
1	150.8	8.0	two weeks later						Average	146.0
2	153.3	8.0	bare		No0 Encapsuated
2	148.2	9.0	bare		Test	T_low (ns)	T_hi (ns)	f_hi (MHz)	f_lo (MHz)	F_mid (MHz)
2	147.4	9.0	hand below antenna		1	6.47	7.22	154.5	138.5	146.5
2	146.0	9.0	hands above and below antenna		2	6.79	7.05	147.2	141.8	144.5
2	148.3	9.0	one week later		3	6.73	7.14	148.5	140.0	144.2
2	148.6	9.0	assembled		4	6.73	7.13	148.5	140.3	144.4
2	149.3	9.0	new tuning circuit						Average	144.9
3	151.5	8.0	bare
3	145.8	9.0	bare
3	146.1	9.0	one week later
3	145.9	9.0	assembled		ID	f1 (MHZ)	f2 (MHz)	f2-f1 (MHz)	f3 (MHz)	f3-f2 (MHz)	f4 (MHz)	f4-f3 (MHz)
3	150.0	8.0	new tuning circuit		1	151.4	152.6	1.2	151.8	-0.8	151.6	-0.2
4	154.4	8.0	bare		2	149.3	151.7	2.4	148.9	-2.8	145.9	-3.0
4	149.5	9.0	bare		3	150.0	152.5	2.5	149.7	-2.8	147.1	-2.6
4	144.6	10.0	bare		4	148.6	149.8	1.2	147.8	-2.0	145.1	-2.7
4	144.6	10.0	one week later		5	148.6	150.0	1.4	147.5	-2.5	145.3	-2.2
4	145.1	10.0	assembled		7	146.7	147.7	1.0	148.5	0.8	147.2	-1.3
4	148.6	9.0	new tuning circuit
5	151.7	8.0	bare			C14	C15	f_bare	f_epoxy
5	145.3	9.0	bare		ID	(pF)	(pF)	(MHz)	(MHz)
5	146.1	9.0	one week later		2.1	9.5	1.0	145
5	146.0	9.0	assembled		2.2	9.0	1.0	147
5	148.6	9.0	new tuning circuit		2.3	9.5	1.0	146
6	152.9	8.0	bare		2.4	9.5	1.0	145
6	147.1	9.0	bare		2.5	9.5	1.0	146
6	147.3	9.0	assembled with power		2.6	10.0	1.0	144
6	143.7	9.0	add liquid epoxy		2.7	9.0	1.0	148
6	146.3	9.0	after epoxy bake
6	146.8	9.0	after three days
6	146.6	9.0	with three coats silicone
6	145.4	9.0	five coats silicone fully cured
7	144.4	9.0	water-soluble flux on parts
7	147.3	9.0	bare
7	147.3	9.0	one week later
7	146.7	9.0	assembled

## Rad_Field
		EZ290	EZ500	EZ290	EZ500	EZ290	EZ500	EZ290	EZ500
	Angle	Intensity	Intensity	relative	relative	power	power	fraction	fraction
0.000	0.0	6.5	6.5	1.000	1.000	0.239	0.239	0.058	0.061
0.196	11.3	5.4	6.4	0.837	0.985	0.393	0.462	0.095	0.117
0.393	22.5	6.1	6.0	0.938	0.923	0.639	0.629	0.155	0.159
0.589	33.8	5.4	5.4	0.837	0.831	0.726	0.720	0.176	0.183
0.785	45.0	4.6	4.5	0.713	0.692	0.727	0.706	0.176	0.179
0.982	56.3	3.7	3.4	0.574	0.523	0.650	0.592	0.158	0.150
1.178	67.5	2.6	2.2	0.403	0.338	0.485	0.407	0.118	0.103
1.374	78.8	1.4	1.0	0.217	0.154	0.266	0.189	0.065	0.048
1.571	90.0	0.0	0.0	0.000	0.000	0.000	0.000	0.000	0.000
						4.124	3.943

## Li_Ion
Discharge					Charge					Discharge					Discharge					Discharge
	A	B					C	D				E	F					G					H
	20 mA-hr	160 mA-hr		20 mA-hr	160 mA-hr			20 mA-hr	160 mA-hr				20 mA-hr	160 mA-hr		20 mA-hr	160 mA-hr			160 mA-hr		160 mA-hr			160 mA-hr		160 mA-hr
Time (s)	V1	V2	Time (hr)	capacity	capacity		Time (s)	V1	V2	Time (hr)		Time (s)	V1	V2	Time (hr)	capacity	capacity		Time (s)	V2	Time (hr)	capacity		Time (s)	V2	Time (hr)	capacity
1367006521	4.34	4.3	0.00	0.00	0.00		1367256377	0.98	3.60	0.00		1367266107	4.35	4.33	0.00	0.00	0.00		1367440827	4.21	0.00	0.00		1367609358	4.22	0.00	0.00
1367006558	4.34	4.3	0.01	0.22	1.10		1367256377	0.98	3.60	0.00		1367266122 / Time (hr) / Voltage (V)	4.32	4.33	0.00	0.09	0.45		1367440828	4.21	0.00	0.06		1367609359	4.22	0.00	0.12
1367006560	4.34	4.3	0.01	0.24	1.16		1367256377	1.09	3.60	0.00		1367266123	4.31	4.33	0.00	0.10	0.48		1367440829	4.21	0.00	0.12		1367609360	4.22	0.00	0.23
1367006561	4.34	4.3	0.01	0.24	1.19		1367256377	1.21	3.60	0.00		1367266134	4.31	4.32	0.01	0.16	0.81		1367440837	4.21	0.00	0.58		1367609420	3.96	0.02	6.83
1367006621	4.23	4.23	0.03	0.59	2.96		1367256378	1.33	3.61	0.00		1367266140	4.07	4.28	0.01	0.20	0.99		1367440897	4.06	0.02	3.97		1367609480	3.88	0.03	13.30
1367006681	4.16	4.2	0.04	0.94	4.71		1367256378	1.44	3.60	0.00		1367266141	4.07	4.27	0.01	0.20	1.02		1367440957	4.02	0.04	7.32		1367609540	3.82	0.05	19.67
1367006741	4.12	4.18	0.06	1.28	6.45		1367256378	1.54	3.61	0.00		1367266143	4.06	4.27	0.01	0.21	1.08		1367441017	3.99	0.05	10.64		1367609600	3.78	0.07	25.97
1367006801	4.09	4.17	0.08	1.62	8.19		1367256378	1.64	3.61	0.00		1367266145	4.06	4.27	0.01	0.22	1.14		1367441077	3.97	0.07	13.95		1367609660	3.75	0.08	32.22
1367006819	4.08	4.16	0.08	1.73	8.71		1367256379	1.75	3.61	0.00		1367266155	4.03	4.26	0.01	0.28	1.43		1367441137	3.94	0.09	17.23		1367609720	3.72	0.10	38.42
1367006823	4.07	4.16	0.08	1.75	8.82		1367256379	1.85	3.61	0.00		1367266165	4	4.25	0.02	0.34	1.73		1367441197	3.92	0.10	20.50		1367609780	3.69	0.12	44.57
1367006835	4.07	4.16	0.09	1.82	9.17		1367256379	1.95	3.61	0.00		1367266175	3.98	4.24	0.02	0.39	2.02		1367441257	3.9	0.12	23.75		1367609840	3.66	0.13	50.67
1367006881	4.05	4.15	0.10	2.08	10.49		1367256379	2.05	3.61	0.00		1367266236	3.9	4.22	0.04	0.72	3.81		1367441317	3.88	0.14	26.98		1367609900	3.64	0.15	56.73
1367006885	4.05	4.15	0.10	2.10	10.61		1367256381	2.14	3.61	0.00		1367266295	3.84	4.2	0.05	1.04	5.53		1367441377	3.86	0.15	30.20		1367609960	3.61	0.17	62.75
1367007185	3.95	4.09	0.18	3.74	19.13		1367256381	2.45	3.61	0.00		1367266356	3.79	4.19	0.07	1.36	7.31		1367441437	3.85	0.17	33.41		1367610020	3.59	0.18	68.73
1367007485	3.86	4.04	0.27	5.35	27.55		1367256381	2.47	3.61	0.00		1367266415	3.75	4.17	0.09	1.66	9.02		1367441497	3.83	0.19	36.60		1367610080	3.57	0.20	74.68
1367007785	3.8	3.99	0.35	6.94	35.86		1367256381	2.49	3.61	0.00		1367266475	3.72	4.16	0.10	1.97	10.75		1367441557	3.82	0.20	39.78		1367610140	3.55	0.22	80.60
1367008085	3.74	3.94	0.43	8.49	44.07		1367256382	2.51	3.61	0.00		1367266480	3.71	4.16	0.10	2.00	10.89		1367441617	3.81	0.22	42.96		1367610200	3.52	0.23	86.47
1367008385	3.69	3.91	0.52	10.03	52.21		1367256382	2.53	3.61	0.00		1367266540	3.68	4.15	0.12	2.31	12.62		1367441677	3.79	0.24	46.12		1367610260	3.49	0.25	92.28
1367008685	3.65	3.88	0.60	11.55	60.30		1367256382	2.54	3.61	0.00		1367266600	3.65	4.14	0.14	2.61	14.35		1367441737	3.78	0.25	49.27		1367610320	3.46	0.27	98.05
1367008985	3.62	3.84	0.68	13.06	68.30		1367256384	2.55	3.61	0.00		1367266660	3.62	4.12	0.15	2.91	16.06		1367441797	3.77	0.27	52.41		1367610380	3.43	0.28	103.77
1367009285	3.59	3.81	0.77	14.56	76.23		1367256384	2.59	3.62	0.00		1367266720	3.6	4.11	0.17	3.21	17.78		1367441857	3.75	0.29	55.53		1367610440	3.38	0.30	109.40
1367009585	3.56	3.79	0.85	16.04	84.13		1367256385	2.59	3.62	0.00		1367266780	3.57	4.1	0.19	3.51	19.48		1367441917	3.74	0.30	58.65		1367610500	3.33	0.32	114.95
1367009885	3.51	3.76	0.93	17.50	91.96		1367256385	2.60	3.62	0.00		1367266840	3.55	4.09	0.20	3.81	21.19		1367441977	3.73	0.32	61.76		1367610560	3.27	0.33	120.40
1367010185	3.44	3.74	1.02	18.94	99.76		1367256385	2.59	3.62	0.00		1367266900	3.53	4.08	0.22	4.10	22.89		1367442037	3.72	0.34	64.86		1367610620	3.2	0.35	125.73
1367010485	3.27	3.72	1.10	20.30	107.51		1367256385	2.60	3.62	0.00		1367266960	3.51	4.07	0.24	4.39	24.58		1367442097	3.71	0.35	67.95		1367610680	3.11	0.37	130.92
1367010785	1.64	3.7	1.18	20.98	115.21		1367256386	2.60	3.62	0.00		1367267020	3.49	4.06	0.25	4.68	26.28		1367442157	3.7	0.37	71.03		1367610740	0.06	0.38	131.02
1367011085	0.74	3.69	1.27	21.29	122.90		1367256386	2.60	3.62	0.00		1367267080	3.47	4.05	0.27	4.97	27.96		1367442217	3.69	0.39	74.11		1367610800	0.06	0.40	131.12
1367011385	0.47	3.68	1.35	21.49	130.57		1367256386	2.61	3.62	0.00		1367267140	3.45	4.05	0.29	5.26	29.65		1367442277	3.68	0.40	77.18		1367610860	0.06	0.42	131.22
1367011685	0.34	3.65	1.43	21.63	138.17		1367256386	2.61	3.62	0.00		1367267200	3.42	4.03	0.30	5.55	31.33		1367442337	3.68	0.42	80.24		1367610920	0.06	0.43	131.32
1367011985	0.26	3.62	1.52	21.74	145.71		1367256387	2.61	3.62	0.00		1367267260	3.4	4.02	0.32	5.83	33.01		1367442397	3.67	0.44	83.30
1367012285	0.21	3.56	1.60	21.82	153.13		1367256387	2.62	3.62	0.00		1367267320	3.38	4.01	0.34	6.11	34.68		1367442457	3.66	0.45	86.35
1367012585	0.18	3.38	1.68	21.90	160.17		1367256388	2.62	3.62	0.00		1367267380	3.37	4.01	0.35	6.39	36.35		1367442517	3.65	0.47	89.39
1367012885	0.15	0.06	1.77	21.96	160.30		1367256388	2.63	3.62	0.00		1367267440	3.35	4	0.37	6.67	38.01		1367442577	3.64	0.49	92.43
1367013185	0.14	0.06	1.85	22.02	160.42		1367256388	2.63	3.62	0.00		1367267500	3.33	3.99	0.39	6.95	39.68		1367442637	3.63	0.50	95.45
1367013485	0.13	0.06	1.93	22.07	160.55		1367256388	2.63	3.62	0.00		1367267560	3.3	3.97	0.40	7.22	41.33		1367442697	3.62	0.52	98.47
1367013785	0.13	0.06	2.02	22.13	160.67		1367256389	2.63	3.62	0.00		1367267620	3.29	3.96	0.42	7.50	42.98		1367442757	3.62	0.54	101.48
1367014085	0.12	0.06	2.10	22.18	160.80		1367256389	2.63	3.62	0.00		1367267680	3.27	3.95	0.44	7.77	44.63		1367442817	3.61	0.55	104.49
1367014385	0.12	0.06	2.18	22.23	160.92		1367256389	2.64	3.62	0.00		1367267740	3.24	3.94	0.45	8.04	46.27		1367442877	3.6	0.57	107.49
1367014685	0.11	0.06	2.27	22.27	161.05		1367256389	2.64	3.62	0.00		1367267800	3.22	3.94	0.47	8.31	47.91		1367442937	3.59	0.59	110.48
1367014985	0.11	0.06	2.35	22.32	161.17		1367256390	2.64	3.62	0.00		1367267860	3.2	3.93	0.49	8.57	49.55		1367442997	3.57	0.60	113.46
1367015285	0.11	0.06	2.43	22.36	161.30		1367256390	2.65	3.62	0.00		1367267920	3.18	3.92	0.50	8.84	51.18		1367443057	3.56	0.62	116.43
1367015585	0.11	0.06	2.52	22.41	161.42		1367256399	2.71	3.64	0.01		1367267980	3.15	3.91	0.52	9.10	52.81		1367443117	3.55	0.64	119.38
1367015885	0.11	0.06	2.60	22.46	161.55		1367256409	2.75	3.65	0.01		1367268040	3.13	3.91	0.54	9.36	54.44		1367443177	3.53	0.65	122.33
1367016185	0.11	0.06	2.68	22.50	161.67		1367256420	2.78	3.66	0.01		1367268100	3.11	3.9	0.55	9.62	56.06		1367443237	3.52	0.67	125.26
1367016485	0.11	0.06	2.77	22.55	161.80		1367256429	2.81	3.67	0.01		1367268160	3.08	3.9	0.57	9.88	57.69		1367443297	3.5	0.69	128.18
1367016785	0.11	0.06	2.85	22.59	161.92		1367256439	2.83	3.68	0.02		1367268220	3.06	3.89	0.59	10.13	59.31		1367443357	3.47	0.70	131.07
1367017085	0.11	0.06	2.93	22.64	162.05		1367256449	2.87	3.69	0.02		1367268280	3.03	3.88	0.60	10.39	60.93		1367443417	3.45	0.72	133.94
1367017385	0.11	0.06	3.02	22.69	162.17		1367256459	2.91	3.70	0.02		1367268340	3.01	3.88	0.62	10.64	62.54		1367443477	3.42	0.74	136.79
				1367256469	2.96	3.70	0.03		1367268400	2.98	3.87	0.64	10.89	64.16		1367443537	3.38	0.75	139.61
				1367256479	3.02	3.71	0.03		1367268460	2.96	3.86	0.65	11.13	65.76		1367443597	3.33	0.77	142.38
				1367256489	3.09	3.72	0.03		1367268520	2.93	3.86	0.67	11.38	67.37		1367443657	3.25	0.79	145.09
				1367256499	3.17	3.73	0.03		1367268580	2.91	3.85	0.69	11.62	68.98		1367443717	3.14	0.80	147.71
				1367256509	3.26	3.74	0.04		1367268640	2.89	3.85	0.70	11.86	70.58		1367443777	0.06	0.82	147.76
				1367256529	3.39	3.75	0.04		1367268700	2.86	3.84	0.72	12.10	72.18		1367443837	0.06	0.84	147.81
				1367256539	3.44	3.76	0.05		1367268760	2.84	3.83	0.74	12.33	73.78		1367443897	0.06	0.85	147.86
				1367256549	3.49	3.76	0.05		1367268820	2.82	3.83	0.75	12.57	75.37		1367443957	0.06	0.87	147.91
				1367256559	3.53	3.77	0.05		1367268880	2.79	3.82	0.77	12.80	76.96		1367444017	0.06	0.89	147.96
				1367256569	3.57	3.77	0.05		1367268940	2.77	3.82	0.79	13.03	78.56		1367444077	0.06	0.90	148.01
				1367256579	3.61	3.78	0.06		1367269000	2.75	3.81	0.80	13.26	80.14		1367444137	0.06	0.92	148.06
				1367256589	3.64	3.78	0.06		1367269060	2.73	3.81	0.82	13.49	81.73		1367444197	0.06	0.94	148.11
				1367256599	3.67	3.79	0.06		1367269120	2.7	3.8	0.84	13.71	83.31		1367444257	0.06	0.95	148.16
				1367256609	3.69	3.79	0.06		1367269180	2.67	3.8	0.85	13.94	84.90		1367444317	0.06	0.97	148.21
				1367256619	3.72	3.80	0.07		1367269240	2.65	3.79	0.87	14.16	86.48		1367444377	0.06	0.99	148.26
				1367256629	3.74	3.80	0.07		1367269300	2.62	3.79	0.89	14.38	88.06
				1367256639	3.76	3.81	0.07		1367269360	2.58	3.78	0.90	14.59	89.63
				1367256649	3.77	3.81	0.08		1367269420	2.54	3.78	0.92	14.80	91.21
				1367256659	3.79	3.81	0.08		1367269480	2.49	3.77	0.94	15.01	92.78
				1367256669	3.81	3.82	0.08		1367269540	2.43	3.77	0.95	15.21	94.35
				1367256679	3.82	3.82	0.08		1367269600	2.34	3.76	0.97	15.41	95.91
				1367256689	3.84	3.82	0.09		1367269660	2.21	3.76	0.99	15.59	97.48
				1367256699	3.85	3.82	0.09		1367269720	2.01	3.76	1.00	15.76	99.05
				1367256709	3.87	3.82	0.09		1367269780	1.72	3.75	1.02	15.90	100.61
				1367256769	3.93	3.83	0.11		1367269840	1.43	3.75	1.04	16.02	102.17
				1367256829	3.98	3.83	0.13		1367269900	1.23	3.74	1.05	16.12	103.73
				1367256889	4.02	3.83	0.14		1367269960	1.07	3.74	1.07	16.21	105.29
				1367256949	4.05	3.84	0.16		1367270020	0.96	3.74	1.09	16.29	106.85
				1367257009	4.08	3.84	0.18		1367270080	0.86	3.73	1.10	16.37	108.40
				1367257069	4.10	3.85	0.19		1367270140	0.78	3.73	1.12	16.43	109.96
				1367257129	4.12	3.86	0.21		1367270200	0.71	3.73	1.14	16.49	111.51
				1367257189	4.14	3.86	0.23		1367270260	0.65	3.72	1.15	16.54	113.06
				1367257249	4.16	3.87	0.24		1367270320	0.6	3.72	1.17	16.59	114.61
				1367257309	4.18	3.87	0.26		1367270380	0.55	3.72	1.19	16.64	116.16
				1367257369	4.19	3.88	0.28		1367270440	0.51	3.72	1.20	16.68	117.71
				1367257429	4.20	3.89	0.29		1367270562	0.47	3.71	1.24	16.76	120.85
				1367257489	4.22	3.89	0.31		1367270571	0.46	3.71	1.24	16.77	121.08
				1367257549	4.23	3.90	0.33		1367270631	0.43	3.71	1.26	16.80	122.63
				1367257609	4.24	3.90	0.34		1367270691	0.4	3.7	1.27	16.84	124.17
				1367257669	4.25	3.91	0.36		1367270707	0.39	3.7	1.28	16.84	124.58
				1367257729	4.26	3.91	0.38		1367270767	0.37	3.7	1.29	16.88	126.12
				1367257789	4.27	3.92	0.39		1367270827	0.34	3.7	1.31	16.90	127.67
				1367257849	4.27	3.92	0.41		1367270887	0.32	3.7	1.33	16.93	129.21
				1367257909	4.28	3.92	0.43		1367270947	0.3	3.69	1.34	16.96	130.75
				1367257969	4.29	3.93	0.44		1367271007	0.28	3.69	1.36	16.98	132.28
				1367258029	4.20	3.93	0.46		1367271067	0.26	3.69	1.38	17.00	133.82
				1367258089	4.18	3.93	0.48		1367271127	0.25	3.68	1.39	17.02	135.35
				1367258149	4.30	3.94	0.49		1367271187	0.23	3.68	1.41	17.04	136.89
				1367258209	4.34	3.94	0.51		1367271247	0.22	3.68	1.43	17.06	138.42
				1367258269	4.35	3.94	0.53		1367271307	0.21	3.67	1.44	17.08	139.95
				1367258329	4.35	3.94	0.54		1367271367	0.2	3.67	1.46	17.09	141.48
				1367258389	4.35	3.95	0.56		1367271427	0.19	3.66	1.48	17.11	143.00
				1367258449	4.35	3.95	0.58		1367271487	0.18	3.66	1.49	17.12	144.53
				1367258509	4.35	3.95	0.59		1367271547	0.17	3.65	1.51	17.14	146.05
				1367258569	4.35	3.95	0.61		1367271607	0.16	3.64	1.53	17.15	147.57
				1367258629	4.35	3.96	0.63		1367271667	0.16	3.64	1.54	17.16	149.08
				1367258689	4.35	3.96	0.64		1367271727	0.15	3.63	1.56	17.18	150.60
				1367258749	4.35	3.96	0.66		1367271787	0.14	3.62	1.58	17.19	152.10
				1367258809	4.34	3.96	0.68		1367271847	0.14	3.61	1.59	17.20	153.61
				1367258869	4.34	3.97	0.69		1367271907	0.14	3.6	1.61	17.21	155.11
				1367258929	4.35	3.97	0.71		1367271967	0.13	3.58	1.63	17.22	156.60
				1367258989	4.35	3.97	0.73		1367272027	0.13	3.56	1.64	17.23	158.08
				1367259049	4.35	3.98	0.74		1367272087	0.13	3.54	1.66	17.24	159.56
				1367259109	4.35	3.98	0.76		1367272147	0.12	3.51	1.68	17.25	161.02
				1367259169	4.35	3.98	0.78		1367272207	0.12	3.47	1.69	17.26	162.47
				1367259229	4.35	3.99	0.79		1367272267	0.12	3.41	1.71	17.27	163.89
				1367259289	4.35	3.99	0.81		1367272327	0.12	3.35	1.73	17.28	165.28
				1367259349	4.35	3.99	0.83		1367272387	0.12	3.27	1.74	17.29	166.65
				1367259409	4.35	4.00	0.84		1367272447	0.12	3.16	1.76	17.30	167.96
				1367259469	4.35	4.00	0.86		1367272507	0.12	0.06	1.78	17.31	167.99
				1367259529	4.35	4.00	0.88		1367272567	0.11	0.06	1.79	17.32	168.01
				1367259589	4.35	4.01	0.89		1367272627	0.11	0.06	1.81	17.33	168.04
				1367259649	4.35	4.01	0.91		1367272687	0.11	0.06	1.83	17.34	168.06
				1367259709	4.35	4.01	0.93		1367272747	0.11	0.06	1.84	17.35	168.09
				1367259769	4.35	4.02	0.94		1367272807	0.11	0.06	1.86	17.36	168.11
				1367259829	4.35	4.02	0.96		1367272867	0.11	0.06	1.88	17.37	168.14
				1367259889	4.34	4.02	0.98		1367272927	0.11	0.06	1.89	17.38	168.16
				1367259949	4.35	4.02	0.99		1367272987	0.11	0.06	1.91	17.39	168.19
				1367260009	4.35	4.03	1.01		1367273047	0.11	0.06	1.93	17.40	168.21
				1367260069	4.35	4.03	1.03		1367273107	0.11	0.06	1.94	17.41	168.24
				1367260129	4.35	4.04	1.04		1367273167	0.11	0.06	1.96	17.42	168.26
				1367260189	4.35	4.04	1.06		1367273227	0.11	0.06	1.98	17.42	168.29
				1367260249	4.35	4.05	1.08		1367273287	0.1	0.06	1.99	17.43	168.31
				1367260309	4.35	4.05	1.09		1367273347	0.11	0.06	2.01	17.44	168.34
				1367260369	4.35	4.05	1.11		1367273407	0.1	0.06	2.03	17.45	168.36
				1367260429	4.35	4.06	1.13		1367273467	0.1	0.06	2.04	17.46	168.39
				1367260489	4.34	4.06	1.14		1367273527	0.11	0.06	2.06	17.47	168.41
				1367260549	4.35	4.07	1.16		1367273587	0.1	0.06	2.08	17.48	168.44
				1367260609	4.35	4.07	1.18		1367273647	0.1	0.06	2.09	17.48	168.46
				1367260669	4.35	4.08	1.19		1367273707	0.1	0.06	2.11	17.49	168.49
				1367260729	4.35	4.08	1.21		1367273767	0.1	0.06	2.13	17.50	168.51
				1367260789	4.35	4.09	1.23		1367273827	0.1	0.06	2.14	17.51	168.54
				1367260849	4.35	4.09	1.24		1367273887	0.1	0.05	2.16	17.52	168.56
				1367260909	4.35	4.10	1.26		1367273947	0.1	0.06	2.18	17.53	168.58
				1367260969	4.35	4.11	1.28		1367274007	0.1	0.06	2.19	17.53	168.61
				1367261029	4.35	4.11	1.29		1367274067	0.1	0.06	2.21	17.54	168.63
				1367261089	4.35	4.11	1.31		1367274127	0.1	0.06	2.23	17.55	168.66
				1367261150	4.35	4.12	1.33		1367274187	0.1	0.06	2.24	17.56	168.68
				1367261209	4.35	4.12	1.34		1367274247	0.1	0.06	2.26	17.57	168.71
				1367261270	4.35	4.13	1.36		1367274307	0.1	0.06	2.28	17.58	168.73
				1367261329	4.35	4.14	1.38		1367274367	0.1	0.06	2.29	17.58	168.76
				1367261389	4.35	4.14	1.39		1367274427	0.1	0.05	2.31	17.59	168.78
				1367261449	4.35	4.15	1.41		1367274487	0.1	0.06	2.33	17.60	168.80
				1367261509	4.35	4.15	1.43		1367274547	0.1	0.06	2.34	17.61	168.83
				1367261569	4.35	4.16	1.44		1367274607	0.1	0.06	2.36	17.62	168.85
				1367261629	4.35	4.17	1.46		1367274667	0.1	0.06	2.38	17.63	168.88
				1367261689	4.35	4.18	1.48		1367274727	0.1	0.06	2.39	17.63	168.90
				1367261749	4.35	4.19	1.49		1367274787	0.1	0.06	2.41	17.64	168.93
				1367261809	4.35	4.19	1.51		1367274847	0.1	0.06	2.43	17.65	168.95
				1367261869	4.35	4.20	1.53		1367274907	0.1	0.06	2.44	17.66	168.98
				1367261929	4.35	4.21	1.54		1367274967	0.1	0.06	2.46	17.67	169.00
				1367261989	4.35	4.22	1.56		1367275027	0.09	0.06	2.48	17.68	169.03
				1367262049	4.35	4.23	1.58		1367275087	0.09	0.06	2.49	17.68	169.05
				1367262109	4.35	4.24	1.59		1367275147	0.09	0.06	2.51	17.69	169.08
				1367262169	4.35	4.25	1.61		1367275207	0.09	0.06	2.53	17.70	169.10
				1367262229	4.35	4.25	1.63		1367275267	0.09	0.06	2.54	17.71	169.13
				1367262289	4.35	4.25	1.64		1367275327	0.09	0.06	2.56	17.71	169.15
				1367262349	4.35	4.26	1.66		1367275387	0.09	0.06	2.58	17.72	169.18
				1367262409	4.35	4.26	1.68		1367275447	0.09	0.06	2.59	17.73	169.20
				1367262469	4.35	4.26	1.69		1367275507	0.09	0.06	2.61	17.74	169.23
				1367262529	4.35	4.26	1.71		1367275567	0.09	0.06	2.63	17.74	169.25
				1367262589	4.35	4.26	1.73		1367275627	0.09	0.06	2.64	17.75	169.28
				1367262649	4.35	4.26	1.74		1367275687	0.09	0.06	2.66	17.76	169.30
				1367262709	4.35	4.26	1.76		1367275747	0.08	0.06	2.68	17.76	169.33
				1367262769	4.35	4.30	1.78		1367275807	0.09	0.06	2.69	17.77	169.35
				1367262829	4.35	4.30	1.79		1367275867	0.08	0.06	2.71	17.78	169.38
				1367262889	4.35	4.31	1.81		1367275927	0.08	0.06	2.73	17.79	169.40
				1367262949	4.35	4.31	1.83		1367275987	0.08	0.06	2.74	17.79	169.43
				1367263009	4.35	4.31	1.84		1367276047	0.08	0.06	2.76	17.80	169.45
				1367263069	4.35	4.31	1.86		1367276107	0.08	0.06	2.78	17.81	169.48
				1367263129	4.35	4.31	1.88		1367276167	0.08	0.06	2.79	17.81	169.50
				1367263189	4.35	4.31	1.89		1367276227	0.08	0.06	2.81	17.82	169.53
				1367263249	4.35	4.32	1.91		1367276287	0.08	0.06	2.83	17.83	169.55
				1367263309	4.35	4.32	1.93		1367276347	0.08	0.06	2.84	17.83	169.58
				1367263369	4.35	4.31	1.94		1367276407	0.08	0.06	2.86	17.84	169.60
				1367263429	4.35	4.32	1.96		1367276467	0.08	0.06	2.88	17.85	169.63
				1367263489	4.35	4.32	1.98		1367276527	0.08	0.06	2.89	17.85	169.65
				1367263549	4.35	4.32	1.99		1367276587	0.08	0.06	2.91	17.86	169.68
				1367263609	4.35	4.32	2.01		1367276647	0.07	0.06	2.93	17.86	169.70
				1367263669	4.35	4.32	2.03		1367276707	0.08	0.06	2.94	17.87	169.73
				1367263729	4.35	4.32	2.04		1367276767	0.08	0.06	2.96	17.88	169.75
				1367263789	4.35	4.32	2.06		1367276827	0.07	0.06	2.98	17.88	169.78
				1367263849	4.35	4.32	2.08		1367276887	0.07	0.06	2.99	17.89	169.80
				1367263909	4.35	4.32	2.09		1367276947	0.07	0.06	3.01	17.90	169.83
				1367263969	4.35	4.33	2.11		1367277007	0.07	0.06	3.03	17.90	169.85
				1367264029	4.35	4.32	2.13		1367277067	0.07	0.06	3.04	17.91	169.88
				1367264089	4.35	4.32	2.14		1367277127	0.07	0.06	3.06	17.91	169.90
				1367264149	4.35	4.33	2.16		1367277187	0.07	0.06	3.08	17.92	169.93
				1367264209	4.35	4.32	2.18		1367277247	0.07	0.05	3.09	17.92	169.95
				1367264269	4.35	4.33	2.19		1367277307	0.07	0.06	3.11	17.93	169.97
				1367264329	4.35	4.33	2.21		1367277367	0.07	0.06	3.13	17.94	170.00
				1367264389	4.35	4.33	2.23		1367277427	0.07	0.06	3.14	17.94	170.02
				1367264449	4.35	4.33	2.24		1367277487	0.07	0.06	3.16	17.95	170.05
				1367264509	4.35	4.33	2.26		1367277547	0.07	0.06	3.18	17.95	170.07
				1367264569	4.35	4.33	2.28		1367277607	0.07	0.06	3.19	17.96	170.10
				1367264629	4.35	4.33	2.29		1367277667	0.07	0.06	3.21	17.97	170.12
				1367264689	4.35	4.33	2.31		1367277727	0.07	0.06	3.23	17.97	170.15
				1367264749	4.35	4.33	2.33		1367277787	0.07	0.06	3.24	17.98	170.17
				1367264809	4.35	4.33	2.34		1367277847	0.07	0.06	3.26	17.98	170.20
				1367264869	4.35	4.33	2.36		1367277907	0.07	0.06	3.28	17.99	170.22
				1367264929	4.35	4.33	2.38		1367277967	0.06	0.06	3.29	17.99	170.25
				1367264989	4.35	4.33	2.39		1367278027	0.06	0.05	3.31	18.00	170.27
				1367265049	4.35	4.33	2.41		1367278087	0.07	0.06	3.33	18.00	170.30
				1367265109	4.35	4.33	2.43		1367278147	0.06	0.06	3.34	18.01	170.32
				1367265169	4.35	4.33	2.44		1367278207	0.06	0.06	3.36	18.01	170.35
				1367265229	4.35	4.33	2.46		1367278267	0.06	0.06	3.38	18.02	170.37
				1367265289	4.35	4.33	2.48		1367278327	0.06	0.06	3.39	18.02	170.40
				1367265349	4.35	4.33	2.49		1367278387	0.06	0.06	3.41	18.03	170.42
				1367265411	4.35	4.33	2.51		1367278447	0.06	0.06	3.43	18.03	170.45
				1367265469	4.35	4.33	2.53		1367278507	0.06	0.06	3.44	18.04	170.47
				1367265529	4.35	4.33	2.54		1367278567	0.06	0.06	3.46	18.04	170.50
				1367265589	4.35	4.33	2.56		1367278627	0.06	0.06	3.48	18.05	170.52
				1367265649	4.35	4.33	2.58		1367278687	0.06	0.05	3.49	18.05	170.54

## Fiber
Fiber Outer Diameter Measurements			Distance (mm) / Diameter (µm)

NA=0.66 Fiber Resolution			300-mm Length
Diameter			Position	Diameter
(µm)			(mm)	(µm)
420			0	420	400	420	410	410	380	420	380	390
420			20	410	410	420	410	410	380	420	380	390
420			40	410	410	420	410	410	380	420	390	390
410			60	410	410	420	400	400	390	420	390	380
420			80	410	420	420	400	400	390	430	390	380
420			100	400	420	420	400	400	400	430	390	380
420			120	400	420	420	400	390	400	440	400	370
420			140	400	420	420	400	390	410	440	400	380
420			160	400	410	410	400	390	410	430	410	380
420			180	400		410	400	380	410	430	410	390
420			200	400		410	390	380	410	430	410	390
420			220	390					410			400
420			240	390					410			410
420			260					410
2.7			280
			300

NA=0.66 Sections		stdev	11.99	µm	max	370
Diameter		average	405.38	µm	min	440
(µm)
420	420	400	380	390	410	400	410	420	420	420	400	400	400	410
400	400	420	410	400	400	400	410	400	400	390	400	410	400	400
410	410	400	400	410	420	390	400	400	370	400	410	410	400	400
410	410	400	400	390	410	410	420	400	410	410	400	390	400	410
400	390	420	400	400	400	410	400	420	390	420	410	390	410	390
400	420	400	400	430	390	400	380	410	400	380	430	420	410	410
410	380	410	400	400	410	410	400	420	400	390	440	400	420	390
400	410	400	400	400	410	400	390	400	420	420	410	400	430	380
420	410	430	410	410	430	410	410	410	400	400	410	400	410	410
420	410	400	400	430	420	400	410	410	400	400	410	410	390	400
410	400	400	410 / Diameter (µm) / Frequency	410	400	390	420	410	390	390	390	400	390	400
400	420	380	430	420	420	390	400	410	400	410	420	410	410	390
420	430	400	400	410	400	400	430	400	430	410	400	390	420	390
410	400	400	400	390	400	410	400	410	420	390	380	400	420	390
390	410	420	400	440	420	430	420	410	400	400	390	390	410	410

value	frequency
360	0
370	1
380	7
390	29
400	81
410	62
420	32
430	11
440	2
450	0
460	0
	225

NA=0.86 Fiber Sections			410.6481481481
Diameter
(µm)
420	410	420	420	420	410	410	430	400	420	420	400	390	420	400
420	410	420	410	390	400	400	420	400	400	410	410	400	410	410
410	400	410	420	420	420	400	420	410	410	410	430	400	310	400
390	410	410	400	410	400	390	420	410	430	420	130	420	130	400
420	420	410	400	430	400	390	420	400	410	420	410	380	410	400
420	400	410	410	400	420	390	410	400	430	400	400	400	400	400
400	420	400	410	440	420	420	410	410	140	420	410	250	420	420
410	390	410	390	410	420	400	410	420	420	410	400	410	400	410
400	400	420	420	410	430	430	400	410	160	410	210	410	410	400
410	390	410	410	410	430	420	410	420	140	420	420	410	130	150
400	400	430	410	410	410	430	410	400	130	400	410	410	190
420	410	430	420	410	410	410	400	410	390	420	400	390	410
			Diameter (µm) / Frequency
value	frequency
390	11
400	42
410	59
420	41
430	11
440	1
450	0
	165

## EZ500_Map
	92	98	100	101	104	104	106	108	110	110	110	111	112	114	113	111	112	112	109	109	108	105	104	101	100	96	86	63
	105	114	118	120	122	123	126	126	126	126	129	130	128	127	129	128	128	127	125	122	122	118	117	112	108	104	93	68
	116	125	128	129	133	135	134	137	138	139	141	140	141	141	140	140	139	137	135	133	131	129	125	120	116	110	99	72		Outer	Core	Optimal		Max	NA=0.66	NA=0.86	NA=0.22
	121	131	137	139	140	139	143	142	148	148	147	149	144	146	147	147	146	144	143	137	138	135	132	127	121	115	103	73		Diameter	Diameter	Location		Incidence	Coupling	Coupling	Coupling
	125	137	142	146	147	146	148	149	152	151	152	151	153	151	151	151	151	148	147	144	142	140	136	131	127	119	105	77		(um)	(µm)	x (µm)	y (µm)	(mW/mW)	(mW/mW)	(mW/mW)	(mW/mW)
	128	140	147	150	152	151	147	147	147	149	149	149	146	147	146	145	143	140	137	133	130	141	143	137	130	123	109	80		220	200	25	0	0.18	0.078	0.133	0.009
	128	142	148	153	156	135	108	113	116	117	120	120	119	122	122	123	123	123	124	119	108	131	146	141	135	127	113	81		330	300	40	0	0.40	0.174	0.296	0.019
	128	142	147	153	153	131	114	147	156	158	161	162	161	160	164	164	163	162	161	154	118	132	148	144	139	130	115	83		390	355	40	0	0.53	0.231	0.392	0.025
	124	137	146	147	147	125	115	152	162	169	170	172	171	170	172	169	170	171	169	160	123	131	150	148	140	132	117	85		440	400	30	0	0.63	0.275	0.466	0.030
	103	111	114	116	114	103	97	121	150	170	176	175	176	175	175	175	178	175	172	162	125	131	153	149	145	133	120	87		484	440	10	0	0.73	0.318	0.540	0.035
	72	75	77	78	80	80	81	92	137	170	176	174	175	176	177	179	175	176	174	167	130	131	155	152	145	136	122	89		511	465	0	0	0.80	0.349	0.592	0.038
	65	68	70	72	76	76	77	88	132	171	176	179	179	177	175	180	177	178	173	168	131	132	156	153	146	136	122	91					Fiber Diameter (µm) / Efficiency (mW/mW)
	63	65	68	69	72	74	76	84	127	170	177	180	177	180	177	177	178	180	174	170	133	130	156	156	147	138	122	92
	61	63	67	72	73	73	75	82	125	169	176	178	179	179	178	179	180	178	176	170	134	131	158	155	148	136	123	94
	60	65	68	77	75	73	75	81	125	168	179	180	178	182	179	182	176	179	176	173	138	131	158	155	149	139	125	94		Outer	Incidence	NA=0.66	NA=0.86
	61	66	69	72	70	73	75	84	122	165	180	181	182	183	181	182	182	179	179	174	138	129	157	154	149	140	126	96		Diameter	Max	Coupling	Coupling
	66	66	69	71	73	73	76	84	123	169	179	182	182	182	182	179	180	182	178	172	141	129	157	156	149	140	126	97		(µm)	(mW/mW)	(mW/mW)	(mW/mW)
	68	70	74	76	76	78	80	88	125	169	180	182	180	181	181	180	181	180	177	173	143	128	155	156	146	138	126	97		300	0.360	0.157	0.266
	91	100	107	111	113	108	98	125	151	171	177	181	180	180	181	179	181	181	176	171	143	125	156	154	149	137	124	98		310	0.372	0.162	0.275
	113	129	137	142	142	133	108	146	163	173	173	175	177	178	176	176	176	177	172	166	140	125	157	152	145	136	125	99		320	0.403	0.176	0.298
	117	136	144	149	150	140	113	135	149	151	157	155	155	154	152	153	150	148	146	141	125	123	154	150	143	138	124	100		330	0.415	0.181	0.307
	118	138	146	149	153	151	122	121	126	131	134	137	137	138	140	144	143	142	140	140	138	142	150	150	144	136	125	101		340	0.453	0.198	0.335
	117	136	147	148	154	154	150	153	158	159	161	161	161	164	163	164	165	160	162	158	154	151	152	147	141	133	123	100		350	0.475	0.207	0.352
	116	135	142	148	154	156	153	157	160	162	161	162	164	165	164	163	164	161	160	156	154	151	145	142	136	130	120	99		360	0.475	0.207	0.352
	112	131	139	145	148	151	152	157	157	158	160	159	159	159	160	160	159	160	157	153	149	146	141	137	132	125	118	99		370	0.515	0.225	0.381
	107	127	133	141	142	147	147	148	151	150	153	153	153	154	154	152	151	151	148	147	143	138	135	132	125	120	111	95		380	0.533	0.232	0.394
	100	119	126	131	135	135	137	138	142	141	141	141	141	141	142	140	139	138	136	133	129	128	126	120	116	110	104	90		390	0.541	0.236	0.400
	92	106	111	114	116	117	118	118	118	119	119	119	118	117	114	115	114	113	110	106	104	102	100	96	93	89	84	75		400	0.572	0.249	0.423
																														410	0.582	0.254	0.431
	44	50	52	53	56	56	58	60	62	62	62	63	64	66	65	63	64	64	61	61	60	57	56	53	52	48	38	15		420	0.624	0.272	0.462
	57	66	70	72	74	75	78	78	78	78	81	82	80	79	81	80	80	79	77	74	74	70	69	64	60	56	45	20		430	0.634	0.276	0.469
	68	77	80	81	85	87	86	89	90	91	93	92	93	93	92	92	91	89	87	85	83	81	77	72	68	62	51	24		440	0.642	0.280	0.475
	73	83	89	91	92	91	95	94	100	100	99	101	96	98	99	99	98	96	95	89	90	87	84	79	73	67	55	25		450	0.690	0.301	0.511
	77	89	94	98	99	98	100	101	104	103	104	103	105	103	103	103	103	100	99	96	94	92	88	83	79	71	57	29		460	0.698	0.304	0.517
	80	92	99	102	104	103	99	99	99	101	101	101	98	99	98	97	95	92	89	85	82	93	95	89	82	75	61	32		470	0.722	0.315	0.534
	80	94	100	105	108	87	60	65	68	69	72	72	71	74	74	75	75	75	76	71	60	83	98	93	87	79	65	33		480	0.763	0.333	0.565
	80	94	99	105	105	83	66	99	108	110	113	114	113	112	116	116	115	114	113	106	70	84	100	96	91	82	67	35		490	0.770	0.336	0.570
	76	89	98	99	99	77	67	104	114	121	122	124	123	122	124	121	122	123	121	112	75	83	102	100	92	84	69	37		500	0.814	0.355	0.602
	55	63	66	68	66	55	49	73	102	122	128	127	128	127	127	127	130	127	124	114	77	83	105	101	97	85	72	39
	24	27	29	30	32	32	33	44	89	122	128	126	127	128	129	131	127	128	126	119	82	83	107	104	97	88	74	41
	17	20	22	24	28	28	29	40	84	123	128	131	131	129	127	132	129	130	125	120	83	84	108	105	98	88	74	43
	15	17	20	21	24	26	28	36	79	122	129	132	129	132	129	129	130	132	126	122	85	82	108	108	99	90	74	44
	13	15	19	24	25	25	27	34	77	121	128	130	131	131	130	131	132	130	128	122	86	83	110	107	100	88	75	46
	12	17	20	29	27	25	27	33	77	120	131	132	130	134	131	134	128	131	128	125	90	83	110	107	101	91	77	46
	13	18	21	24	22	25	27	36	74	117	132	133	134	135	133	134	134	131	131	126	90	81	109	106	101	92	78	48
	18	18	21	23	25	25	28	36	75	121	131	134	134	134	134	131	132	134	130	124	93	81	109	108	101	92	78	49
	20	22	26	28	28	30	32	40	77	121	132	134	132	133	133	132	133	132	129	125	95	80	107	108	98	90	78	49
	43	52	59	63	65	60	50	77	103	123	129	133	132	132	133	131	133	133	128	123	95	77	108	106	101	89	76	50
	65	81	89	94	94	85	60	98	115	125	125	127	129	130	128	128	128	129	124	118	92	77	109	104	97	88	77	51
	69	88	96	101	102	92	65	87	101	103	109	107	107	106	104	105	102	100	98	93	77	75	106	102	95	90	76	52
	70	90	98	101	105	103	74	73	78	83	86	89	89	90	92	96	95	94	92	92	90	94	102	102	96	88	77	53
	69	88	99	100	106	106	102	105	110	111	113	113	113	116	115	116	117	112	114	110	106	103	104	99	93	85	75	52
	68	87	94	100	106	108	105	109	112	114	113	114	116	117	116	115	116	113	112	108	106	103	97	94	88	82	72	51
	64	83	91	97	100	103	104	109	109	110	112	111	111	111	112	112	111	112	109	105	101	98	93	89	84	77	70	51
	59	79	85	93	94	99	99	100	103	102	105	105	105	106	106	104	103	103	100	99	95	90	87	84	77	72	63	47
	52	71	78	83	87	87	89	90	94	93	93	93	93	93	94	92	91	90	88	85	81	80	78	72	68	62	56	42
	44	58	63	66	68	69	70	70	70	71	71	71	70	69	66	67	66	65	62	58	56	54	52	48	45	41	36	27

	-230	-213	-196	-179	-162	-145	-128	-111	-94	-77	-60	-43	-26	-9	8	25	42	59	76	93	110	127	144	161	178	195	212	229
-230	0.13	0.14	0.15	0.17	0.18	0.20	0.22	0.23	0.25	0.26	0.27	0.27	0.27	0.27	0.27	0.27	0.27	0.26	0.25	0.24	0.22	0.21	0.20	0.18	0.16	0.15	0.13	0.12
-213	0.13	0.15	0.17	0.18	0.20	0.22	0.24	0.25	0.27	0.29	0.30	0.31	0.31	0.31	0.31	0.31	0.30	0.29	0.28	0.27	0.25	0.24	0.22	0.20	0.19	0.17	0.15	0.13
-196	0.14	0.16	0.18	0.20	0.21	0.23	0.25	0.27	0.29	0.31	0.33	0.34	0.34	0.34	0.34	0.34	0.34	0.33	0.32	0.30	0.28	0.26	0.25	0.23	0.21	0.19	0.17	0.15
-179	0.15	0.17	0.19	0.21	0.23	0.25	0.27	0.29	0.32	0.34	0.35	0.36	0.37	0.37	0.38	0.38	0.37	0.36	0.35	0.33	0.31	0.29	0.27	0.25	0.23	0.21	0.18	0.16
-162	0.16	0.18	0.20	0.22	0.24	0.27	0.29	0.32	0.34	0.36	0.38	0.39	0.40	0.41	0.41	0.41	0.40	0.40	0.38	0.36	0.34	0.32	0.30	0.27	0.25	0.23	0.20	0.18
-145	0.16	0.19	0.21	0.23	0.26	0.28	0.31	0.33	0.36	0.39	0.41	0.42	0.43	0.43	0.44	0.44	0.44	0.43	0.41	0.39	0.37	0.34	0.32	0.30	0.27	0.24	0.22	0.19
-128	0.17	0.19	0.22	0.25	0.27	0.30	0.33	0.36	0.38	0.41	0.43	0.45	0.45	0.46	0.47	0.47	0.47	0.46	0.44	0.42	0.40	0.37	0.34	0.32	0.29	0.26	0.24	0.21
-111	0.18	0.20	0.23	0.26	0.29	0.32	0.35	0.38	0.41	0.43	0.45	0.47	0.48	0.49	0.49	0.50	0.50	0.49	0.47	0.45	0.42	0.40	0.37	0.34	0.31	0.28	0.25	0.22
-94	0.19	0.22	0.24	0.28	0.31	0.34	0.37	0.40	0.43	0.45	0.47	0.49	0.50	0.51	0.52	0.52	0.52	0.52	0.50	0.48	0.45	0.42	0.39	0.36	0.33	0.30	0.27	0.24
-77	0.19	0.23	0.26	0.29	0.32	0.36	0.39	0.42	0.45	0.47	0.49	0.51	0.52	0.53	0.54	0.54	0.54	0.54	0.52	0.50	0.47	0.44	0.41	0.38	0.35	0.31	0.28	0.25
-60	0.20	0.24	0.27	0.31	0.34	0.37	0.41	0.44	0.47	0.49	0.51	0.52	0.53	0.54	0.55	0.56	0.56	0.55	0.54	0.52	0.49	0.46	0.43	0.39	0.36	0.33	0.29	0.26
-43	0.21	0.25	0.28	0.32	0.35	0.38	0.42	0.45	0.48	0.50	0.52	0.53	0.54	0.55	0.56	0.56	0.57	0.57	0.55	0.53	0.51	0.47	0.44	0.41	0.37	0.33	0.30	0.26
-26	0.21	0.25	0.28	0.32	0.36	0.39	0.42	0.45	0.48	0.51	0.52	0.53	0.54	0.55	0.56	0.57	0.57	0.57	0.56	0.54	0.51	0.48	0.45	0.41	0.37	0.34	0.30	0.27
-9	0.21	0.25	0.29	0.32	0.36	0.39	0.43	0.46	0.49	0.51	0.52	0.53	0.54	0.55	0.56	0.57	0.58	0.58	0.57	0.55	0.52	0.49	0.45	0.42	0.38	0.34	0.31	0.27
8	0.21	0.25	0.29	0.33	0.36	0.40	0.43	0.46	0.49	0.51	0.53	0.54	0.55	0.56	0.57	0.58	0.58	0.58	0.57	0.55	0.52	0.49	0.45	0.42	0.38	0.35	0.31	0.27
25	0.21	0.25	0.29	0.32	0.36	0.40	0.43	0.46	0.49	0.51	0.53	0.54	0.55	0.56	0.57	0.58	0.58	0.58	0.57	0.55	0.52	0.49	0.45	0.42	0.38	0.35	0.31	0.27
42	0.21	0.25	0.29	0.32	0.36	0.39	0.43	0.46	0.49	0.51	0.53	0.54	0.55	0.56	0.57	0.58	0.58	0.58	0.57	0.55	0.52	0.49	0.45	0.42	0.38	0.34	0.31	0.27
59	0.21	0.24	0.28	0.32	0.35	0.39	0.42	0.45	0.48	0.50	0.52	0.54	0.55	0.56	0.57	0.58	0.58	0.57	0.56	0.54	0.51	0.48	0.44	0.41	0.37	0.34	0.30	0.27
76	0.20	0.23	0.27	0.30	0.34	0.37	0.40	0.43	0.46	0.49	0.51	0.53	0.54	0.55	0.56	0.56	0.57	0.56	0.55	0.52	0.49	0.46	0.43	0.40	0.36	0.33	0.29	0.26
93	0.19	0.22	0.25	0.29	0.32	0.35	0.38	0.41	0.44	0.47	0.49	0.51	0.52	0.53	0.54	0.54	0.54	0.54	0.52	0.50	0.47	0.44	0.41	0.38	0.35	0.31	0.28	0.25
110	0.18	0.21	0.24	0.27	0.30	0.33	0.36	0.39	0.42	0.45	0.47	0.49	0.50	0.51	0.52	0.52	0.52	0.51	0.50	0.48	0.45	0.42	0.39	0.36	0.33	0.30	0.27	0.24
127	0.18	0.20	0.23	0.26	0.28	0.31	0.34	0.37	0.40	0.43	0.45	0.47	0.48	0.49	0.49	0.49	0.49	0.49	0.47	0.45	0.42	0.39	0.37	0.34	0.31	0.28	0.25	0.22
144	0.17	0.19	0.22	0.24	0.27	0.30	0.32	0.35	0.38	0.40	0.43	0.44	0.45	0.46	0.46	0.47	0.46	0.46	0.44	0.42	0.39	0.37	0.34	0.32	0.29	0.26	0.24	0.21
161	0.16	0.18	0.21	0.23	0.26	0.28	0.31	0.33	0.36	0.38	0.40	0.42	0.42	0.43	0.43	0.43	0.43	0.42	0.41	0.39	0.37	0.34	0.32	0.30	0.27	0.25	0.22	0.19
178	0.16	0.18	0.20	0.22	0.24	0.26	0.29	0.31	0.33	0.36	0.38	0.39	0.40	0.40	0.40	0.40	0.40	0.39	0.38	0.36	0.34	0.32	0.29	0.27	0.25	0.23	0.20	0.18
195	0.15	0.17	0.19	0.21	0.23	0.25	0.27	0.29	0.31	0.33	0.35	0.36	0.36	0.37	0.37	0.37	0.37	0.36	0.34	0.33	0.31	0.29	0.27	0.25	0.23	0.21	0.18	0.16
212	0.14	0.16	0.17	0.19	0.21	0.23	0.25	0.27	0.29	0.30	0.32	0.33	0.33	0.33	0.34	0.33	0.33	0.32	0.31	0.30	0.28	0.26	0.24	0.22	0.21	0.19	0.17	0.15
229	0.13	0.15	0.16	0.18	0.19	0.21	0.23	0.25	0.26	0.28	0.29	0.29	0.30	0.30	0.30	0.30	0.29	0.29	0.28	0.26	0.25	0.23	0.22	0.20	0.18	0.17	0.15	0.13
246	0.12	0.13	0.15	0.16	0.18	0.19	0.21	0.22	0.24	0.25	0.26	0.26	0.26	0.26	0.26	0.26	0.26	0.25	0.24	0.23	0.22	0.21	0.19	0.18	0.16	0.15	0.13	0.11

## LED_Power
	Photodiode Sensitivity at 460 nm				0.18	mA/mW

Head	Power (mW) at Current			Photcurrent (mA) at Current				2.9
Fixture	30 mA	50 mA	70 mA	30 mA	50 mA	70 mA		Current	Photocurrent	Power
2.1	27.5	41.2	51.7	4.95	7.42	9.30		mA	mA	mW
2.2	25.7	38.9	51.7	4.62	7.00	9.31		0	0	0.00
2.3	27.1	40.1	49.4	4.88	7.22	8.89		1.68	0.04	0.50
2.4	26.7	40.1	55.3	4.81	7.22	9.95		5.76	0.13	1.63
2.5	26.9	40.7	50.1	4.85	7.32	9.02		12.75	0.35	4.38
2.6	26.8	41.1	55.3	4.82	7.40	9.96		18.43	0.47	5.88
2.7	27.6	41.8	51.7	4.96	7.53	9.30		22.22	0.59	7.38
2.8	26.3	40.0	54.8	4.74	7.20	9.87		25.39	0.69	8.63
2.9	28.2	42.1	51.9	5.07	7.58	9.34		29.2	0.86	10.75
average	27.0	40.7	52.4	4.9	7.3	9.4		30.65	0.89	11.13
stdev	0.7	1.0	2.2	0.1	0.2	0.4		38.3	1.11	13.88
								44.6	1.26	15.75
Angle (deg) / Intensity (mW/cm2)								51.2	1.47	18.38
Angle	Ip	Power						63.7	1.64	20.50
(deg)	(uW)	(mW/cm2)						74.4	1.81	22.63
90	21	0.12						87.4	2.01	25.13
60	88	0.49						Forward Current (mA) / Taper Power (mW)
45	100	0.56
0	102	0.57
-45	90	0.50
-60	63	0.35
-90	11	0.06

C460EZ500	Photodiode Sensitivity at 460 nm				0.18	mA/mW
Head	Power (mW) at Current			Diameter	Coupling
Fixture	30 mA	50 mA	70 mA	(um)	(%)	Comment
2.1	27.5	41.2	51.7	450	0.0	Broke bond wire during gluing
2.2	25.7	38.9	51.7	440	22.2	Off-center by 200 um
2.3	27.1	40.1	49.4	440	35.0	Off-center by 100 um
2.4	26.7	40.1	55.3	440	41.2	Off-center by 40 um
2.5	26.9	40.7	50.1	440	38.6	Well-centered
2.6	26.8	41.1	55.3	460	38.9	Off-center by 100 um
2.7	27.6	41.8	51.7	460	37.7	Well-centered, tacky glue
2.8	26.3	40.0	54.8	440	0.0	Broke fiber after glue set
2.9	28.2	42.1	51.9	440	38.3	Well-centered
2.10	26.7			440	38.2	Well-centered
average	26.9	40.7	52.4
stdev	0.7	1.0	2.2

C527EZ500	Photodiode Sensitivity at 527 nm				0.23	mA/mW
Head	Power (mW) at Current			Diameter	Coupling
Fixture	30 mA	50 mA	70 mA	(um)	(%)	Comment
2.11	13.3	20.3	25.0	440	38.8	Well-centered
2.12	13.2	19.5	26.0	440	39.2	Well-centered
2.13	13.3	19.1	23.2
2.14	12.3	18.4	24.0
2.15	13.8	20.3	25.2
average	13.2	19.5	24.7
stdev	0.5	0.8	1.1

Head		Guide	Power (mW) @ 40mA		Coupling
Fixture	Color	Cannula	LED	Taper	(%)
6.13	Blue	Steel	34	15	44
6.14	Blue	Silica	36	16	44
6.15	Blue	Silica	35	15	43
6.16	Green	Silica	15	5.0	33
6.17	Blue	Silica	34	13	38
6.18	Blue	Silica	34	15	44
6.19	Green	Silica	16	6.7	42
6.20	Green	Silica	16	6.5	41
6.22	Blue	Silica	33	14	42
6.23	Blue	Steel	35	18	51
6.5	Blue	Silica	35	16	46
6.6	Blue	Steel	34	15	44
6.7	Blue	Silica	35	15	43
6.8	Blue	Silica	34	15	44
7.1	Blue	Silica	35	15	43
				Average	42.9
Note: 7.1 is A3024HFD				Stdev	3.9

## IQ
VBAT = 3.6 V
	U1 Switch	Current (uA)			Current (mA)		Efficiency (%)		Lead
Lamp	Threshold	No U2	No U2	All	Lamp 100%	Lamp 10%	Lamp 100%	Lamp 10%	Resistance
Circuit	(mV)	0 Hz	100 Hz	0 Hz	@ 42 mA	@ 42 mA	@ 42 mA	@ 42 mA	(Ohm)
2.1	7	0.6	0.8	0.6	65.3	8.6	87.2	66.2	56
2.2	7	0.6	0.8	2.6	65.2	8.5	87.3	67.0	55
2.3	5	0.6	0.9	0.7	62.6	8.3	91.0	68.6	55
2.4	5	0.6	0.8	2.3	63.6	8.6	89.5	66.2	51
2.5	4	0.7	0.8	2.5	64.0	8.8	89.0	64.7	58
2.6	6	0.7	1.0	2.3	63.5	8.4	89.7	67.8	56
2.7	6	0.6	0.8	0.8	63.8	8.4	89.3	67.8	57
Ave	5.7	0.6	0.8	1.7	64.0	8.5	89.0	66.9	55.4

A: 32mA 100% S			Time (min) / Voltage (V)
time	VBAT
(min)	(V)
1	4.2
2	4.0
3	3.9
4	3.8
8	3.6
10	3.5
12	3.3

B: 32mA 10% S
time	VBAT
(min)	(V)
1	4.14
3	4.11
5	4.09
12	4.06
22	4.01
31	3.97
37	3.95
39	3.94
43	3.93				D: 42mA 100% L
46	3.91				time	VBAT
67	3.83				(min)	(V)
81	3.79				0.1	4.15
85	3.78				2	4.1
108	3.72				9	4.04
114	3.71				14	4.02
129	3.69				17	4.01
136	3.68				24	3.95
160	3.63				29	3.91
165	3.61				36	3.87
171	3.58				45	3.83
185	3.45				49	3.83
188	3.41				93	3.66
					100	3.65
C: 42mA 10% S					105	3.63
time	VBAT				110	3.61
(min)	(V)				113	3.59
0.1	4.2				120	3.52
10	4.03				124	3.31
12	4.02
19	3.98
28	3.93
36	3.89
43	3.87
50	3.83
52	3.82
55	3.81
62	3.78
65	3.77
77	3.73
90	3.7
97	3.68
108	3.66
115	3.64
120	3.63
124	3.61
126	3.6
130	3.57
133	3.55
136	3.53
140	3.47
142	3.45